## Sheet1
Attribute Name | Tag | Type | Attribute Description | Source
Image Type | (0008,0008) | 1 | Image identification characteristics. See Section C.8.2.1.1.1 for specialization. | Table C.8-3. CT Image Module Attributes
Samples per Pixel | (0028,0002) | 1 | Number of samples (planes) in this image. See Section C.8.2.1.1.2 for specialization. | Table C.8-3. CT Image Module Attributes
Photometric Interpretation | (0028,0004) | 1 | Specifies the intended interpretation of the pixel data. See Section C.8.2.1.1.3 for specialization. | Table C.8-3. CT Image Module Attributes
Bits Allocated | (0028,0100) | 1 | Number of bits allocated for each pixel sample. Each sample shall have the same number of bits allocated. See Section C.8.2.1.1.4 for specialization. | Table C.8-3. CT Image Module Attributes
Bits Stored | (0028,0101) | 1 | Number of bits stored for each pixel sample. Each sample shall have the same number of bits stored. See Section C.8.2.1.1.5 for specialization. | Table C.8-3. CT Image Module Attributes
High Bit | (0028,0102) | 1 | Most significant bit for pixel sample data. Each sample shall have the same high bit. See Section C.8.2.1.1.6 for specialization. | Table C.8-3. CT Image Module Attributes
Rescale Intercept | (0028,1052) | 1 | The value b in relationship between stored values (SV) and the output units. Output units = m*SV+b If Image Type (0008,0008) Value 1 is ORIGINAL and Value 3 is not LOCALIZER, output units shall be Hounsfield Units (HU). | Table C.8-3. CT Image Module Attributes
Rescale Slope | (0028,1053) | 1 | m in the equation specified in Rescale Intercept (0028,1052). | Table C.8-3. CT Image Module Attributes
Rescale Type | (0028,1054) | 1C | Specifies the output units of Rescale Slope (0028,1053) and Rescale Intercept (0028,1052). See Section C.11.1.1.2 for Defined Terms and further explanation. Required if the Rescale Type is not HU (Hounsfield Units). May be present otherwise. | Table C.8-3. CT Image Module Attributes
KVP | (0018,0060) | 2 | Peak kilo voltage output of the x-ray generator used | Table C.8-3. CT Image Module Attributes
Acquisition Number | (0020,0012) | 2 | A number identifying the single continuous gathering of data over a period of time that resulted in this image | Table C.8-3. CT Image Module Attributes
Scan Options | (0018,0022) | 3 | Parameters of scanning sequence. | Table C.8-3. CT Image Module Attributes
Data Collection Diameter | (0018,0090) | 3 | The diameter in mm of the region over which data were collected | Table C.8-3. CT Image Module Attributes
Data Collection Center (Patient) | (0018,9313) | 3 | The x, y, and z coordinates (in the patient coordinate system) in mm of the center of the region in which data were collected. See Section C.8.15.3.6.1. | Table C.8-3. CT Image Module Attributes
Reconstruction Diameter | (0018,1100) | 3 | Diameter in mm of the region from within which data were used in creating the reconstruction of the image. Data may exist outside this region and portions of the patient may exist outside this region. | Table C.8-3. CT Image Module Attributes
Reconstruction Target Center (Patient) | (0018,9318) | 3 | The x, y, and z coordinates (in the patient coordinate system) of the reconstruction center target point as used for reconstruction in mm. See Section C.8.15.3.6.1. | Table C.8-3. CT Image Module Attributes
Distance Source to Detector | (0018,1110) | 3 | Distance in mm from source to detector center. This value is traditionally referred to as Source Image Receptor Distance (SID). | Table C.8-3. CT Image Module Attributes
Distance Source to Patient | (0018,1111) | 3 | Distance in mm from source to isocenter (center of field of view). This value is traditionally referred to as Source Object Distance (SOD). | Table C.8-3. CT Image Module Attributes
Gantry/Detector Tilt | (0018,1120) | 3 | Nominal angle of tilt in degrees of the scanning gantry. Not intended for mathematical computations. | Table C.8-3. CT Image Module Attributes
Table Height | (0018,1130) | 3 | The distance in mm of the top of the patient table to the center of rotation; below the center is positive. | Table C.8-3. CT Image Module Attributes
Rotation Direction | (0018,1140) | 3 | Direction of rotation of the source when relevant, about nearest principal axis of equipment. Enumerated Values: CW (clockwise), CC (counter clockwise) | Table C.8-3. CT Image Module Attributes
Exposure Time | (0018,1150) | 3 | Time of x-ray exposure in msec. If Acquisition Type (0018,9302) equals SPIRAL, the value of this attribute shall be Revolution Time (0018,9305) divided by the Spiral Pitch Factor (0018,9311). See Section C.8.15.3.8.1 and Section C.8.15.3.2.1. | Table C.8-3. CT Image Module Attributes
X-Ray Tube Current | (0018,1151) | 3 | X-Ray Tube Current in mA. | Table C.8-3. CT Image Module Attributes
Exposure | (0018,1152) | 3 | The exposure expressed in mAs, for example calculated from Exposure Time and X-Ray Tube Current. | Table C.8-3. CT Image Module Attributes
Exposure in ÂµAs | (0018,1153) | 3 | The exposure expressed in ÂµAs, for example calculated from Exposure Time and X-Ray Tube Current. | Table C.8-3. CT Image Module Attributes
Filter Type | (0018,1160) | 3 | Label for the type of filter inserted into the x-ray beam. | Table C.8-3. CT Image Module Attributes
Generator Power | (0018,1170) | 3 | Power in kW to the x-ray generator. | Table C.8-3. CT Image Module Attributes
Focal Spot(s) | (0018,1190) | 3 | Size of the focal spot in mm. For devices with variable focal spot or multiple focal spots, small dimension followed by large dimension. | Table C.8-3. CT Image Module Attributes
Convolution Kernel | (0018,1210) | 3 | A label describing the convolution kernel or algorithm used to reconstruct the data | Table C.8-3. CT Image Module Attributes
Revolution Time | (0018,9305) | 3 | The time in seconds of a complete revolution of the source around the gantry orbit. | Table C.8-3. CT Image Module Attributes
Single Collimation Width | (0018,9306) | 3 | The width of a single row of acquired data (in mm). Adjacent physical detector rows may have been combined to form a single effective acquisition row. | Table C.8-3. CT Image Module Attributes
Total Collimation Width | (0018,9307) | 3 | The width of the total collimation (in mm) over the area of active x-ray detection. This will be equal the number of effective detector rows multiplied by single collimation width. | Table C.8-3. CT Image Module Attributes
Table Speed | (0018,9309) | 3 | The distance in mm that the table moves in one second during the gathering of data that resulted in this image. | Table C.8-3. CT Image Module Attributes
Table Feed per Rotation | (0018,9310) | 3 | Motion of the table (in mm) during a complete revolution of the source around the gantry orbit. | Table C.8-3. CT Image Module Attributes
Spiral Pitch Factor | (0018,9311) | 3 | Ratio of the Table Feed per Rotation (0018,9310) to the Total Collimation Width (0018,9307). | Table C.8-3. CT Image Module Attributes
Exposure Modulation Type | (0018,9323) | 3 | A label describing the type of exposure modulation used for the purpose of limiting the dose. Defined Terms: NONE | Table C.8-3. CT Image Module Attributes
Estimated Dose Saving | (0018,9324) | 3 | A percent value of dose saving due to the use of Exposure Modulation Type (0018,9323). A negative percent value of dose savings reflects an increase of exposure. | Table C.8-3. CT Image Module Attributes
CTDIvol | (0018,9345) | 3 | Computed Tomography Dose Index (CTDIvol), in mGy according to IEC 60601-2-44, Ed.2.1 (Clause 29.1.103.4), The Volume CTDIvol. It describes the average dose for this image for the selected CT conditions of operation. | Table C.8-3. CT Image Module Attributes
CTDI Phantom Type Code Sequence | (0018,9346) | 3 | The type of phantom used for CTDI measurement according to IEC 60601-2-44. Only a single Item is permitted in this Sequence. | Table C.8-3. CT Image Module Attributes
Calcium Scoring Mass Factor Patient | (0018,9351) | 3 | The calibration factor for the calcium mass score. These factors incorporate the effects of KV value of the CT image, the patient size, machine specific corrections. See Section C.8.2.1.1.7. | Table C.8-3. CT Image Module Attributes
Calcium Scoring Mass Factor Device | (0018,9352) | 3 | The calibration factors for the calcium mass score of the device. These factors incorporate the effects of KV value of the CT image, machine specific corrections. This a multi-value attribute, the first value specifies the mass factor for a small patient size, the second value for a medium patient size and the third value for a large patient size. See Section C.8.2.1.1.7. | Table C.8-3. CT Image Module Attributes
Energy Weighting Factor | (0018,9353) | 1C | The weighting factor of the data from the primary source in a multiple energy composition image. This factor incorporates the effects of the specific X-Ray source and kV value, examination specific characteristics. Required if one Derivation Code Sequence (0008,9215) Item value is (113097, DCM, \Multi-energy proportional weighting\"). May be present otherwise." | Table C.8-3. CT Image Module Attributes
CT Additional X-Ray Source Sequence | (0018,9360) | 3 | Contains the attributes describing additional acquisition parameters beyond the primary source in a multiple X-Ray source system or a multi-energy acquisition. The primary X-Ray source is specified in other attributes of this module. One or more Items are permitted in this sequence. | Table C.8-3. CT Image Module Attributes
X-Ray Tube Current in mA | (0018,9330) | 1 | Nominal X-Ray tube current in milliamperes. | Table C.8-3. CT Image Module Attributes
Filter Material | (0018,7050) | 1 | The X-Ray absorbing material used in the filter. | Table C.8-3. CT Image Module Attributes
Exposure in mAs | (0018,9332) | 3 | The exposure expressed in milliampere seconds, for example calculated from exposure time and X-Ray tube current. | Table C.8-3. CT Image Module Attributes
Isocenter Position | (300A,012C) | 3 | Isocenter coordinates (x,y,z), in mm. Specifies the location of the machine isocenter in the patient-based coordinate system associated with the Frame of Reference. It allows transformation from the equipment-based coordinate system to the patient-based coordinate system. | Table C.8-3. CT Image Module Attributes
Anatomic Region Sequence | (0008,2218) | 3 | Sequence that identifies the anatomic region of interest in this Instance (i.e., external anatomy, surface anatomy, or general region of the body). | Table 10-7 General Anatomy Optional Macro Attributes
Anatomic Region Modifier Sequence | (0008,2220) | 3 | Sequence of Items that modifies the anatomic region of interest of this Instance. | Table 10-7 General Anatomy Optional Macro Attributes
Primary Anatomic Structure Sequence | (0008,2228) | 3 | Sequence of Items that identifies the primary anatomic structure(s) of interest in this Instance. | Table 10-8 Primary Anatomic Structure Macro Attributes
Primary Anatomic Structure Modifier Sequence | (0008,2230) | 3 | Sequence of Items that modifies the primary anatomic structure of interest in this Instance. | Table 10-8 Primary Anatomic Structure Macro Attributes
View Code Sequence | (0054,0220) | 3 | Sequence that describes the projection of the anatomic region of interest. | Table 10-25 Optional View and Slice Progression Direction Attributes
View Modifier Code Sequence | (0054,0222) | 3 | View Modifier. | Table 10-25 Optional View and Slice Progression Direction Attributes
Slice Progression Direction | (0054,0500) | 3 | Describes the anatomical direction in which a set of slices is progressing. | Table 10-25 Optional View and Slice Progression Direction Attributes
Patient Support Angle | (300A,0122) | 3 | Patient Support angle, i.e., orientation of IEC PATIENT SUPPORT coordinate system with respect to IEC FIXED REFERENCE coordinate system (degrees). | Table 10-27. RT Equipment Correlation Macro Attributes Description
Table Top Pitch Angle | (300A,0140) | 3 | Table Top Pitch Angle, i.e., the rotation of the IEC TABLE TOP coordinate system about the X-axis of the IEC TABLE TOP coordinate system (degrees). See Section C.8.8.25.6.2. | Table 10-27. RT Equipment Correlation Macro Attributes Description
Table Top Roll Angle | (300A,0144) | 3 | Table Top Roll Angle, i.e., the rotation of the IEC TABLE TOP coordinate system about the Y-axis of the IEC TABLE TOP coordinate system (degrees). See Section C.8.8.25.6.2. | Table 10-27. RT Equipment Correlation Macro Attributes Description
Table Top Longitudinal Position | (300A,0129) | 3 | Table Top Longitudinal position in IEC TABLE TOP coordinate system (mm). | Table 10-27. RT Equipment Correlation Macro Attributes Description
Table Top Lateral Position | (300A,012A) | 3 | Table Top Lateral position in IEC TABLE TOP coordinate system (mm). | Table 10-27. RT Equipment Correlation Macro Attributes Description
Image Type | (0008,0008) | 1 | Image identification characteristics. See Section C.8.3.1.1.1 for specialization. | Table C.8-4. MR Image Module Attributes
Samples per Pixel | (0028,0002) | 1 | Number of samples (planes) in this image. See Section C.8.3.1.1.2 for specialization. | Table C.8-4. MR Image Module Attributes
Photometric Interpretation | (0028,0004) | 1 | Specifies the intended interpretation of the pixel data. See Section C.8.3.1.1.3 for specialization. | Table C.8-4. MR Image Module Attributes
Bits Allocated | (0028,0100) | 1 | Number of bits allocated for each pixel sample. Each sample shall have the same number of bits allocated. See Section C.8.3.1.1.4 for specialization. | Table C.8-4. MR Image Module Attributes
Scanning Sequence | (0018,0020) | 1 | Description of the type of data taken. | Table C.8-4. MR Image Module Attributes
Sequence Variant | (0018,0021) | 1 | Variant of the Scanning Sequence. | Table C.8-4. MR Image Module Attributes
Scan Options | (0018,0022) | 2 | Parameters of scanning sequence. | Table C.8-4. MR Image Module Attributes
MR Acquisition Type | (0018,0023) | 2 | Identification of data encoding scheme. | Table C.8-4. MR Image Module Attributes
Repetition Time | (0018,0080) | 2C | The period of time in msec between the beginning of a pulse sequence and the beginning of the succeeding (essentially identical) pulse sequence. | Table C.8-4. MR Image Module Attributes
Echo Time | (0018,0081) | 2 | Time in ms between the middle of the excitation pulse and the peak of the echo produced (kx=0). | Table C.8-4. MR Image Module Attributes
Echo Train Length | (0018,0091) | 2 | Number of lines in k-space acquired per excitation per image. | Table C.8-4. MR Image Module Attributes
Inversion Time | (0018,0082) | 2C | Time in msec after the middle of inverting RF pulse to middle of excitation pulse to detect the amount of longitudinal magnetization. | Table C.8-4. MR Image Module Attributes
Trigger Time | (0018,1060) | 2C | Time, in msec, between peak of the R wave and the peak of the echo produced. | Table C.8-4. MR Image Module Attributes
Sequence Name | (0018,0024) | 3 | User defined name for the combination of Scanning Sequence (0018,0020) and Sequence Variant (0018,0021). | Table C.8-4. MR Image Module Attributes
Angio Flag | (0018,0025) | 3 | Angio Image Indicator. Primary image for Angio processing. | Table C.8-4. MR Image Module Attributes
Number of Averages | (0018,0083) | 3 | Number of times a given pulse sequence is repeated before any parameter is changed | Table C.8-4. MR Image Module Attributes
Imaging Frequency | (0018,0084) | 3 | Precession frequency in MHz of the nucleus being addressed | Table C.8-4. MR Image Module Attributes
Imaged Nucleus | (0018,0085) | 3 | Nucleus that is resonant at the imaging frequency. Examples: 31P, 1H | Table C.8-4. MR Image Module Attributes
Echo Number(s) | (0018,0086) | 3 | The echo number used in generating this image. | Table C.8-4. MR Image Module Attributes
Magnetic Field Strength | (0018,0087) | 3 | Nominal field strength of MR magnet, in Tesla | Table C.8-4. MR Image Module Attributes
Spacing Between Slices | (0018,0088) | 3 | Spacing between slices, in mm. The spacing is measured from the center-to-center of each slice. | Table C.8-4. MR Image Module Attributes
Number of Phase Encoding Steps | (0018,0089) | 3 | Total number of lines in k-space in the 'y' direction collected during acquisition. | Table C.8-4. MR Image Module Attributes
Percent Sampling | (0018,0093) | 3 | Fraction of acquisition matrix lines acquired, expressed as a percent. | Table C.8-4. MR Image Module Attributes
Percent Phase Field of View | (0018,0094) | 3 | Ratio of field of view dimension in phase direction to field of view dimension in frequency direction, expressed as a percent. | Table C.8-4. MR Image Module Attributes
Pixel Bandwidth | (0018,0095) | 3 | Reciprocal of the total sampling period, in hertz per pixel. | Table C.8-4. MR Image Module Attributes
Nominal Interval | (0018,1062) | 3 | Average R-R interval used for the scans, in msec | Table C.8-4. MR Image Module Attributes
Beat Rejection Flag | (0018,1080) | 3 | Beat length sorting has been applied. | Table C.8-4. MR Image Module Attributes
Low R-R Value | (0018,1081) | 3 | R-R interval low limit for beat rejection, in msec | Table C.8-4. MR Image Module Attributes
High R-R Value | (0018,1082) | 3 | R-R interval high limit for beat rejection, in msec | Table C.8-4. MR Image Module Attributes
Intervals Acquired | (0018,1083) | 3 | Number of R-R intervals acquired. | Table C.8-4. MR Image Module Attributes
Intervals Rejected | (0018,1084) | 3 | Number of R-R intervals rejected. | Table C.8-4. MR Image Module Attributes
PVC Rejection | (0018,1085) | 3 | Description of type of PVC rejection criteria used. | Table C.8-4. MR Image Module Attributes
Skip Beats | (0018,1086) | 3 | Number of beats skipped after a detected arrhythmia. | Table C.8-4. MR Image Module Attributes
Heart Rate | (0018,1088) | 3 | Beats per minute. | Table C.8-4. MR Image Module Attributes
Cardiac Number of Images | (0018,1090) | 3 | Number of images per cardiac cycle. | Table C.8-4. MR Image Module Attributes
Trigger Window | (0018,1094) | 3 | Percent of R-R interval, based on Heart Rate (0018,1088), prescribed as a window for a valid/usable trigger. | Table C.8-4. MR Image Module Attributes
Reconstruction Diameter | (0018,1100) | 3 | Diameter in mm. of the region from within which data were used in creating the reconstruction of the image. | Table C.8-4. MR Image Module Attributes
Receive Coil Name | (0018,1250) | 3 | Receive coil used. | Table C.8-4. MR Image Module Attributes
Transmit Coil Name | (0018,1251) | 3 | Transmit coil used. | Table C.8-4. MR Image Module Attributes
Acquisition Matrix | (0018,1310) | 3 | Dimensions of the acquired frequency /phase data before reconstruction. | Table C.8-4. MR Image Module Attributes
In-plane Phase Encoding Direction | (0018,1312) | 3 | The axis of phase encoding with respect to the image. | Table C.8-4. MR Image Module Attributes
Flip Angle | (0018,1314) | 3 | Steady state angle in degrees to which the magnetic vector is flipped from the magnetic vector of the primary field. | Table C.8-4. MR Image Module Attributes
SAR | (0018,1316) | 3 | Calculated whole body Specific Absorption Rate in watts/kilogram. | Table C.8-4. MR Image Module Attributes
Variable Flip Angle Flag | (0018,1315) | 3 | Flip angle variation applied during image acquisition. | Table C.8-4. MR Image Module Attributes
dB/dt | (0018,1318) | 3 | The rate of change of the gradient coil magnetic flux density with time (T/s). | Table C.8-4. MR Image Module Attributes
Temporal Position Identifier | (0020,0100) | 3 | Temporal order of a dynamic or functional set of Images. | Table C.8-4. MR Image Module Attributes
Number of Temporal Positions | (0020,0105) | 3 | Total number of temporal positions prescribed. | Table C.8-4. MR Image Module Attributes
Temporal Resolution | (0020,0110) | 3 | Time delta between Images in a dynamic or functional set of Images. | Table C.8-4. MR Image Module Attributes
Anatomic Region Sequence | (0008,2218) | 3 | Sequence that identifies the anatomic region of interest in this Instance (i.e., external anatomy, surface anatomy, or general region of the body). | Table 10-7 General Anatomy Optional Macro Attributes
Anatomic Region Modifier Sequence | (0008,2220) | 3 | Sequence of Items that modifies the anatomic region of interest of this Instance. | Table 10-7 General Anatomy Optional Macro Attributes
Primary Anatomic Structure Sequence | (0008,2228) | 3 | Sequence of Items that identifies the primary anatomic structure(s) of interest in this Instance. | Table 10-8 Primary Anatomic Structure Macro Attributes
Primary Anatomic Structure Modifier Sequence | (0008,2230) | 3 | Sequence of Items that modifies the primary anatomic structure of interest in this Instance. | Table 10-8 Primary Anatomic Structure Macro Attributes
View Code Sequence | (0054,0220) | 3 | Sequence that describes the projection of the anatomic region of interest. | Table 10-25 Optional View and Slice Progression Direction Attributes
View Modifier Code Sequence | (0054,0222) | 3 | View Modifier. | Table 10-25 Optional View and Slice Progression Direction Attributes
Slice Progression Direction | (0054,0500) | 3 | Describes the anatomical direction in which a set of slices is progressing. | Table 10-25 Optional View and Slice Progression Direction Attributes
Isocenter Position | (300A,012C) | 3 | Isocenter coordinates (x,y,z), in mm. Specifies the location of the machine isocenter in the patient-based coordinate system associated with the Frame of Reference. | Table C.8-4. MR Image Module Attributes
Modality | (0008,0060) | 1 | Type of equipment that originally acquired the data used to create the images in this Series. | Table C.8-68. DX Series Module Attributes
Referenced Performed Procedure Step Sequence | (0008,1111) | 1C | Uniquely identifies the Performed Procedure Step SOP Instance to which the Series is related. Only a single Item shall be included in this sequence. Required if a Performed Procedure Step SOP Class was involved in the creation of this Series. | Table C.8-68. DX Series Module Attributes
Presentation Intent Type | (0008,0068) | 1 | Identifies the intent of the images that are contained within this Series. | Table C.8-68. DX Series Module Attributes
Modality | (0008,0060) | 1 | Type of equipment that originally acquired the data used to create the images in this Series. | Table C.8-73. Mammography Series Module Attributes
Request Attributes Sequence | (0040,0275) | 3 | Sequence that contains attributes from the Imaging Service Request. One or more Items are permitted in this sequence. | Table C.8-73. Mammography Series Module Attributes
Image Laterality | (0020,0062) | 1 | Laterality of (possibly paired) body part (as described in Anatomic Region Sequence (0008,2218)) examined. | Table C.8-69. DX Anatomy Imaged Module Attributes
Anatomic Region Sequence | (0008,2218) | 2 | Sequence that identifies the anatomic region of interest in this Instance (i.e., external anatomy, surface anatomy, or general region of the body). Zero or one Item shall be included in this Sequence. | Table 10-6. General Anatomy Required Macro Attributes
Anatomic Region Modifier Sequence | (0008,2220) | 3 | Sequence of Items that modifies the anatomic region of interest of this Instance. One or more Items are permitted in this sequence. | Table 10-6. General Anatomy Required Macro Attributes
Primary Anatomic Structure Sequence | (0008,2228) | 3 | Sequence of Items that identifies the primary anatomic structure(s) of interest in this Instance. One or more Items are permitted in this sequence. | Table 10-8. Primary Anatomic Structure Macro Attributes
Primary Anatomic Structure Modifier Sequence | (0008,2230) | 3 | Sequence of Items that modifies the primary anatomic structure of interest in this Instance. One or more Items are permitted in this sequence. | Table 10-8. Primary Anatomic Structure Macro Attributes
Image Type | (0008,0008) | 1 | Image identification characteristics. See Section C.8.11.3.1.1 for specialization. | Table C.8-70. DX Image Module Attributes
Samples per Pixel | (0028,0002) | 1 | Number of samples in this image. Enumerated Values: 1 | Table C.8-70. DX Image Module Attributes
Photometric Interpretation | (0028,0004) | 1 | Specifies the intended interpretation of the pixel data. Enumerated Values: MONOCHROME1, MONOCHROME2 | Table C.8-70. DX Image Module Attributes
Bits Allocated | (0028,0100) | 1 | Number of bits allocated for each pixel sample. Enumerated Values: 8, 16 | Table C.8-70. DX Image Module Attributes
Bits Stored | (0028,0101) | 1 | Number of bits stored for each pixel sample. Enumerated Values: 6, 7, 8, 9, 10, 11, 12, 13, 14, 15, 16 | Table C.8-70. DX Image Module Attributes
High Bit | (0028,0102) | 1 | Most significant bit for pixel sample data. Shall be one less than the value in Bit Stored (0028,0101). | Table C.8-70. DX Image Module Attributes
Pixel Representation | (0028,0103) | 1 | Data representation of the pixel samples. Enumerated Values: 0000H (Unsigned Integer) | Table C.8-70. DX Image Module Attributes
Pixel Intensity Relationship | (0028,1040) | 1 | The relationship between the Pixel sample values and the X-Ray beam intensity. Enumerated Values: LIN, LOG. See Section C.8.11.3.1.2 for further explanation. | Table C.8-70. DX Image Module Attributes
Pixel Intensity Relationship Sign | (0028,1041) | 1 | The sign of the relationship between the Pixel sample values stored in Pixel Data (7FE0,0010) and the X-Ray beam intensity. Enumerated Values: 1, -1. See Section C.8.11.3.1.2 for further explanation. | Table C.8-70. DX Image Module Attributes
Rescale Intercept | (0028,1052) | 1 | The value b in the relationship between stored values (SV) in Pixel Data (7FE0,0010) and the output units specified in Rescale Type (0028,1054). Output units = m*SV + b. Enumerated Values: 0. See Section C.8.11.3.1.2 for further explanation. | Table C.8-70. DX Image Module Attributes
Rescale Slope | (0028,1053) | 1 | m in the equation specified by Rescale Intercept (0028,1052). Enumerated Values: 1. See Section C.8.11.3.1.2 for further explanation. | Table C.8-70. DX Image Module Attributes
Rescale Type | (0028,1054) | 1 | Specifies the output units of Rescale Slope (0028,1053) and Rescale Intercept (0028,1052). Enumerated Values: US (Unspecified). See Section C.8.11.3.1.2 for further explanation. | Table C.8-70. DX Image Module Attributes
Presentation LUT Shape | (2050,0020) | 1 | Specifies an identity transformation for the Presentation LUT, other than to account for the value of Photometric Interpretation (0028,0004), such that the output of all grayscale transformations defined in the IOD containing this Module are defined to be P-Values. Enumerated Values: IDENTITY, INVERSE. See Section C.8.11.3.1.2 for further explanation. | Table C.8-70. DX Image Module Attributes
Lossy Image Compression | (0028,2110) | 1 | Specifies whether an Image has undergone lossy compression (at a point in its lifetime). Enumerated Values: 00 (Image has NOT been subjected to lossy compression), 01 (Image has been subjected to lossy compression). Once this value has been set to 01 it shall not be reset. See Section C.7.6.1.1.5. | Table C.8-70. DX Image Module Attributes
Lossy Image Compression Ratio | (0028,2112) | 1C | Describes the approximate lossy compression ratio(s) that have been applied to this image. See Section C.7.6.1.1.5 for further explanation. May be multi-valued if successive lossy compression steps have been applied; the value order shall correspond to the values of Lossy Image Compression Method (0028,2114). Required if Lossy Image Compression (0028,2110) is '01'. | Table C.8-70. DX Image Module Attributes
Derivation Description | (0008,2111) | 3 | A text description of how this image was derived. See Section C.8.11.3.1.4 for further explanation. | Table C.8-70. DX Image Module Attributes
Acquisition Device Processing Description | (0018,1400) | 3 | Indicates any visual processing performed on the images prior to exchange. See Section C.8.11.3.1.3 for further explanation. | Table C.8-70. DX Image Module Attributes
Acquisition Device Processing Code | (0018,1401) | 3 | Code representing the device-specific processing associated with the image (e.g., Organ Filtering code). This Code is manufacturer specific but provides useful annotation information to the knowledgeable observer. | Table C.8-70. DX Image Module Attributes
Patient Orientation | (0020,0020) | 1C | Patient direction of the rows and columns of the image. See Section C.7.6.1.1.1 for further explanation. Required if View Code Sequence (0054,0220) is not present or is present with an Item value other than (G-8300, SRT, 'tissue specimen') or (G-8310, SRT, 'tissue specimen from breast'). May be present otherwise. | Table C.8-70. DX Image Module Attributes
Calibration Image | (0050,0004) | 3 | Indicates whether a reference object (phantom) of known size is present in the image and was used for calibration. Enumerated Values: YES, NO. Device is identified using the Device module. See Section C.7.6.12 for further explanation. | Table C.8-70. DX Image Module Attributes
Burned In Annotation | (0028,0301) | 1 | Indicates whether or not image contains sufficient burned in annotation to identify the patient and date the image was acquired. Enumerated Values: YES, NO. | Table C.8-70. DX Image Module Attributes
VOI LUT Sequence | (0028,3010) | 1C | Defines a sequence of VOI LUTs. See Section C.8.11.3.1.5 for further explanation. One or more Items shall be included in this Sequence. Required if Presentation Intent Type (0008,0068) is FOR PRESENTATION and Window Center (0028,1050) is not present. May also be present if Window Center (0028,1050) is present. | Table C.8-70. DX Image Module Attributes
>LUT Descriptor | (0028,3002) | 1 | Specifies the format of the LUT Data in this Sequence. See Section C.8.11.3.1.5 for further explanation. | Table C.8-70. DX Image Module Attributes
>LUT Explanation | (0028,3003) | 3 | Free form text explanation of the meaning of the LUT. | Table C.8-70. DX Image Module Attributes
>LUT Data | (0028,3006) | 1 | LUT Data in this Sequence. | Table C.8-70. DX Image Module Attributes
Window Center | (0028,1050) | 1C | Defines a Window Center for display. See Section C.8.11.3.1.5 for further explanation. Required if Presentation Intent Type (0008,0068) is FOR PRESENTATION and VOI LUT Sequence (0028,3010) is not present. May also be present if VOI LUT Sequence (0028,3010) is present. | Table C.8-70. DX Image Module Attributes
Window Width | (0028,1051) | 1C | Window Width for display. See Section C.8.11.3.1.5 for further explanation. Required if Window Center (0028,1050) is sent. | Table C.8-70. DX Image Module Attributes
Window Center & Width Explanation | (0028,1055) | 3 | Free form explanation of the meaning of the Window Center and Width. Multiple values correspond to multiple Window Center and Width values. | Table C.8-70. DX Image Module Attributes
Detector Active Time | (0018,7014) | 3 | Time in mSec that the detector is active during acquisition of this image. | Table C.8-71. DX Detector Module Attributes
Detector Activation Offset From Exposure | (0018,7016) | 3 | Offset time in mSec that the detector becomes active after the X-Ray beam is turned on during acquisition of this image. May be negative. | Table C.8-71. DX Detector Module Attributes
Field of View Shape | (0018,1147) | 3 | Shape of the Field of View, that is the image pixels stored in Pixel Data (7FE0,0010). | Table C.8-71. DX Detector Module Attributes
Field of View Dimension(s) | (0018,1149) | 3 | Dimensions in mm of the Field of View, that is the image pixels stored in Pixel Data (7FE0,0010). If Field of View Shape (0018,1147) is: RECTANGLE: row dimension followed by column. ROUND: diameter. HEXAGONAL: diameter of a circumscribed circle. | Table C.8-71. DX Detector Module Attributes
Field of View Origin | (0018,7030) | 1C | Offset of the TLHC of a rectangle circumscribing the Field of View, that is the image pixels stored in Pixel Data (7FE0,0010), before rotation or flipping, from the TLHC of the physical detector area measured in physical detector pixels as a row offset followed by a column offset. Required if Field of View Rotation (0018,7032) or Field of View Horizontal Flip (0018,7034) is present. | Table C.8-71. DX Detector Module Attributes
Field of View Rotation | (0018,7032) | 1C | Clockwise rotation in degrees of Field of View, that is the image pixels stored in Pixel Data (7FE0,0010), relative to the physical detector. Required if Field of View Horizontal Flip (0018,7034) is present. | Table C.8-71. DX Detector Module Attributes
Field of View Horizontal Flip | (0018,7034) | 1C | Whether or not a horizontal flip has been applied to the Field of View, that is the image pixels stored in Pixel Data (7FE0,0010), after rotation relative to the physical detector as described in Field of View Rotation (0018,7032). Required if Field of View Rotation (0018,7032) is present. | Table C.8-71. DX Detector Module Attributes
Imager Pixel Spacing | (0018,1164) | 1 | Physical distance measured at the front plane of the detector housing between the center of each image pixel specified by a numeric pair - row spacing value(delimiter) column spacing value in mm. | Table C.8-71. DX Detector Module Attributes
Cassette ID | (0018,1007) | 3 | Identifier of the cassette that contains the photostimulable phosphor plate, for CR acquisitions. | Table C.8-71. DX Detector Module Attributes
Plate ID | (0018,1004) | 3 | Identifier of the photostimulable phosphor plate, for CR acquisitions. | Table C.8-71. DX Detector Module Attributes
Detector Type | (0018,7004) | 2 | The type of detector used to acquire this image. | Table C.8-71b. Digital X-Ray Detector Macro Attributes
Detector Configuration | (0018,7005) | 3 | The physical configuration of the detector. | Table C.8-71b. Digital X-Ray Detector Macro Attributes
Detector Description | (0018,7006) | 3 | Free text description of detector. | Table C.8-71b. Digital X-Ray Detector Macro Attributes
Detector Mode | (0018,7008) | 3 | Text description of operating mode of detector (implementation specific). | Table C.8-71b. Digital X-Ray Detector Macro Attributes
Detector ID | (0018,700A) | 3 | The ID or serial number of the detector used to acquire this image. | Table C.8-71b. Digital X-Ray Detector Macro Attributes
Date of Last Detector Calibration | (0018,700C) | 3 | The date on which the detector used to acquire this image as identified in Detector ID (0018,700A) was last calibrated. | Table C.8-71b. Digital X-Ray Detector Macro Attributes
Time of Last Detector Calibration | (0018,700E) | 3 | The time at which the detector used to acquire this image as identified in Detector ID (0018,700A) was last calibrated. | Table C.8-71b. Digital X-Ray Detector Macro Attributes
Exposures on Detector Since Last Calibration | (0018,7010) | 3 | Total number of X-Ray exposures that have been made on the detector used to acquire this image as identified in Detector ID (0018,700A) since it was calibrated. | Table C.8-71b. Digital X-Ray Detector Macro Attributes
Exposures on Detector Since Manufactured | (0018,7011) | 3 | Total number of X-Ray exposures that have been made on the detector used to acquire this image as identified in Detector ID (0018,700A) since it was manufactured. | Table C.8-71b. Digital X-Ray Detector Macro Attributes
Detector Time Since Last Exposure | (0018,7012) | 3 | Time in Seconds since an exposure was last made on this detector prior to the acquisition of this image. | Table C.8-71b. Digital X-Ray Detector Macro Attributes
Detector Binning | (0018,701A) | 3 | Number of active detectors used to generate a single pixel. Specified as number of row detectors per pixel then column. | Table C.8-71b. Digital X-Ray Detector Macro Attributes
Detector Manufacturer Name | (0018,702A) | 3 | Name of the manufacturer of the detector component of the acquisition system | Table C.8-71b. Digital X-Ray Detector Macro Attributes
Detector Manufacturer's Model Name | (0018,702B) | 3 | Model name of the detector component of the acquisition system | Table C.8-71b. Digital X-Ray Detector Macro Attributes
Detector Conditions Nominal Flag | (0018,7000) | 3 | Whether or not the detector is operating within normal tolerances during this image acquisition. | Table C.8-71b. Digital X-Ray Detector Macro Attributes
Detector Temperature | (0018,7001) | 3 | Detector temperature during exposure in degrees Celsius. | Table C.8-71b. Digital X-Ray Detector Macro Attributes
Sensitivity | (0018,6000) | 3 | Detector sensitivity in manufacturer specific units. | Table C.8-71b. Digital X-Ray Detector Macro Attributes
Detector Element Physical Size | (0018,7020) | 3 | Physical dimensions of each detector element that comprises the detector matrix, in mm. | Table C.8-71b. Digital X-Ray Detector Macro Attributes
Detector Element Spacing | (0018,7022) | 3 | Physical distance between the center of each detector element, specified by a numeric pair - row spacing value(delimiter) column spacing value in mm. | Table C.8-71b. Digital X-Ray Detector Macro Attributes
Detector Active Shape | (0018,7024) | 3 | Shape of the active area. | Table C.8-71b. Digital X-Ray Detector Macro Attributes
Detector Active Dimension(s) | (0018,7026) | 3 | Dimensions in mm of the active area. | Table C.8-71b. Digital X-Ray Detector Macro Attributes
Detector Active Origin | (0018,7028) | 3 | Offset of the TLHC of a rectangle circumscribing the active detector area from the TLHC of a rectangle circumscribing the physical detector area, measured in physical detector pixels as a row offset followed by a column offset. | Table C.8-71b. Digital X-Ray Detector Macro Attributes
Pixel Spacing | (0028,0030) | 1C | Physical distance in the patient between the center of each pixel, specified by a numeric pair - adjacent row spacing (delimiter) adjacent column spacing in mm. | Table 10-10. Basic Pixel Spacing Calibration Macro Attributes
Pixel Spacing Calibration Type | (0028,0A02) | 3 | The type of correction for the effect of geometric magnification or calibration against an object of known size, if any. | Table 10-10. Basic Pixel Spacing Calibration Macro Attributes
Pixel Spacing Calibration Description | (0028,0A04) | 1C | A free text description of the type of correction or calibration performed. | Table 10-10. Basic Pixel Spacing Calibration Macro Attributes
Exposure Index | (0018,1411) | 3 | Measure of the detector response to radiation in the relevant image region of an image acquired with a digital x-ray imaging system as defined in IEC 62494-1. | Table 10-23. Exposure Index Macro Attributes
Target Exposure Index | (0018,1412) | 3 | The target value used to calculate Deviation Index (0018,1413) as defined in IEC 62494-1. | Table 10-23. Exposure Index Macro Attributes
Deviation Index | (0018,1413) | 3 | A scaled representation of the difference of the Exposure Index compared to the Target Exposure Index as defined in IEC 62494-1 and the report of AAPM TG 116. | Table 10-23. Exposure Index Macro Attributes
Image Type | (0008,0008) | 1 | Image identification characteristics. See Section C.8.11.7.1.4 for specialization. | Table C.8-74. Mammography Image Module Attributes
Positioner Type | (0018,1508) | 1 | Enumerated Values: MAMMOGRAPHIC NONE | Table C.8-74. Mammography Image Module Attributes
Distance Source to Detector | (0018,1110) | 3 | Distance in mm from source to detector center on the chest wall line. See Section C.8.11.7.1.1. | Table C.8-74. Mammography Image Module Attributes
Distance Source to Patient | (0018,1111) | 3 | Distance in mm from source to the breast support side that is closest to the Imaging Subject, as measured along the X-Ray beam vector. See notes for this attribute in Section C.8.11.5 DX Positioning Module. See Section C.8.11.7.1.1 for description of X-Ray beam vector. | Table C.8-74. Mammography Image Module Attributes
Positioner Primary Angle | (0018,1510) | 3 | Signed position in degrees of the X-Ray beam vector in the coronal anatomical plane as if the patient were standing facing the equipment where vertical is zero. | Table C.8-74. Mammography Image Module Attributes
Positioner Primary Angle Direction | (0018,9559) | 3 | The interpretation of the sign of the Positioner Primary Angle (0018,1510). Enumerated Values: CW - Clockwise. Positioner Primary Angle movement is positive when movement is from vertical to the patient's right (with patient standing, facing equipment). CC - Counter clockwise. Positioner Primary Angle is positive when movement is from vertical to the patient's left (with patient standing, facing the equipment). See Section C.8.31.2.1. | Table C.8-74. Mammography Image Module Attributes
Positioner Secondary Angle | (0018,1511) | 3 | Position in degrees of the X-Ray beam vector in the sagittal anatomical plane as if the patient were standing where movement of the X-Ray source from anterior to posterior is positive, and vertical is zero. | Table C.8-74. Mammography Image Module Attributes
Image Laterality | (0020,0062) | 1 | Laterality of the region examined. Enumerated Values: R - right L - left B - both (e.g., cleavage) | Table C.8-74. Mammography Image Module Attributes
Organ Exposed | (0040,0318) | 1 | Organ to which Organ Dose (0040,0316) applies. Enumerated Values: BREAST | Table C.8-74. Mammography Image Module Attributes
Breast Implant Present | (0028,1300) | 3 | Whether or not the imaged breast contains a breast implant regardless of the visibility of a breast implant in the Pixel Data. Enumerated Values: YES NO | Table C.8-74. Mammography Image Module Attributes
Partial View | (0028,1350) | 3 | Indicates whether this image is a partial view, that is a subset of a single view of the breast. Enumerated Values: YES NO | Table C.8-74. Mammography Image Module Attributes
Partial View Description | (0028,1351) | 3 | Free text description of the portion of the breast captured in a partial view image. | Table C.8-74. Mammography Image Module Attributes
Partial View Code Sequence | (0028,1352) | 3 | Sequence that describes the portion or section of the breast captured in a partial view image. See Section C.8.11.7.1.3. | Table C.8-74. Mammography Image Module Attributes
View Code Sequence | (0054,0220) | 1 | Sequence that describes the projection of the anatomic region of interest on the image receptor. Only a single Item shall be included in this Sequence. | Table C.8-74. Mammography Image Module Attributes
View Modifier Code Sequence | (0054,0222) | 2 | View modifier. Zero or more Items shall be included in this Sequence. | Table C.8-74. Mammography Image Module Attributes
Biopsy Target Sequence | (0018,2041) | 3 | Sequence that identifies the targets of a biopsy procedure. One or more Items are permitted in this sequence, where the coordinate system is the same for all Items within a pair of stereo images. | Table C.8-74. Mammography Image Module Attributes
Target UID | (0018,2042) | 1 | Unique identifier for the target. Note To identify the same target in corresponding minus and plus stereo images. | Table C.8-74. Mammography Image Module Attributes
Localizing Cursor Position | (0018,2043) | 1 | Coordinates of localizing cursor position with respect to pixels in the image specified by a column (delimiter) row pair. The values shall be in the range 0\0 (TLHC) to Columns\Rows (BRHC). | Table C.8-74. Mammography Image Module Attributes
Calculated Target Position | (0018,2044) | 1 | The calculated target position (x, y, z) in mm in an equipment relative right-handed coordinate system where the origin is under the central X-Ray beam at the chest wall, and a z value of zero is located at the surface that the breast is in contact with that is closest to the detector. X is parallel to the chest wall, y is positive toward the patient's nipple, and z is positive toward the X-Ray source. | Table C.8-74. Mammography Image Module Attributes
Displayed Z Value | (0018,2046) | 1 | The z value in mm displayed to the user at the time of biopsy. Note This may be the same as the z value of Calculated Target Position or different in direction or reference point, such as relative to the compression paddle. This is not the distance that the needle was inserted, particularly for a lateral (parallel) approach. | Table C.8-74. Mammography Image Module Attributes
Target Label | (0018,2045) | 3 | Target description. | Table C.8-74. Mammography Image Module Attributes
Anatomic Region Sequence | (0008,2218) | 1 | Sequence that identifies the anatomic region of interest in this Instance (i.e., external anatomy, surface anatomy, or general region of the body). Only a single Item shall be included in this sequence. | Table 10-5. General Anatomy Mandatory Macro Attributes
Anatomic Region Modifier Sequence | (0008,2220) | 3 | Sequence of Items that modifies the anatomic region of interest of this Instance. One or more Items are permitted in this sequence. | Table 10-5. General Anatomy Mandatory Macro Attributes
Primary Anatomic Structure Sequence | (0008,2228) | 3 | Sequence of Items that identifies the primary anatomic structure(s) of interest in this Instance. One or more Items are permitted in this sequence. | Table 10-8. Primary Anatomic Structure Macro Attributes
Primary Anatomic Structure Modifier Sequence | (0008,2230) | 3 | Sequence of Items that modifies the primary anatomic structure of interest in this Instance. One or more Items are permitted in this sequence. | Table 10-8. Primary Anatomic Structure Macro Attributes
Body Part Examined | (0018,0015) | 2 | Text description of the part of the body examined. See Annex L “Correspondence of Anatomic Region Codes and Body Part Examined Defined Terms” in PS3.16 for Defined Terms | Table C.8-1. CR Series Module Attributes
View Position | (0018,5101) | 2 | Radiographic view associated with Patient Position (0018,5100). For humans: Defined Terms: AP Anterior/Posterior PA Posterior/Anterior LL Left Lateral RL Right Lateral RLD Right Lateral Decubitus LLD Left Lateral Decubitus RLO Right Lateral Oblique LLO Left Lateral Oblique For animals, capitalized versions with hyphens replaced by underscores of the abbreviations defined in 'Smallwood et al. A Nomenclature for Radiographic Projections Used In Veterinary Medicine' shall be used. See CID 7484 “DX View for Animals”, in which these abbreviations are listed. | Table C.8-1. CR Series Module Attributes
Filter Type | (0018,1160) | 3 | Label for the type of filter inserted into the x-ray beam | Table C.8-1. CR Series Module Attributes
Collimator/grid Name | (0018,1180) | 3 | Label describing any grid inserted. | Table C.8-1. CR Series Module Attributes
Focal Spot(s) | (0018,1190) | 3 | Size of the focal spot in mm. For devices with variable focal spot or multiple focal spots, small dimension followed by large dimension. | Table C.8-1. CR Series Module Attributes
Plate Type | (0018,1260) | 3 | Label of the type of storage phosphor plates used in this series | Table C.8-1. CR Series Module Attributes
Phosphor Type | (0018,1261) | 3 | Label of type of phosphor on the plates | Table C.8-1. CR Series Module Attributes
Photometric Interpretation | (0028,0004) | 1 | Specifies the intended interpretation of the pixel data. | Table C.8-2. CR Image Module Attributes
KVP | (0018,0060) | 3 | Peak kilo voltage output of the x-ray generator used | Table C.8-2. CR Image Module Attributes
Plate ID | (0018,1004) | 3 | The ID or serial number of the sensing plate upon which the image was acquired | Table C.8-2. CR Image Module Attributes
Distance Source to Detector | (0018,1110) | 3 | Distance in mm from source to detector center. | Table C.8-2. CR Image Module Attributes
Distance Source to Patient | (0018,1111) | 3 | Distance in mm from source to center of field of view. | Table C.8-2. CR Image Module Attributes
Exposure Time | (0018,1150) | 3 | Time of x-ray exposure in msec | Table C.8-2. CR Image Module Attributes
X-Ray Tube Current | (0018,1151) | 3 | X-Ray Tube Current in mA. | Table C.8-2. CR Image Module Attributes
Exposure | (0018,1152) | 3 | The exposure expressed in mAs, for example calculated from Exposure Time and X-Ray Tube Current. | Table C.8-2. CR Image Module Attributes
Exposure in ÂµAs | (0018,1153) | 3 | The exposure expressed in ÂµAs, for example calculated from Exposure Time and X-Ray Tube Current. | Table C.8-2. CR Image Module Attributes
Imager Pixel Spacing | (0018,1164) | 3 | Physical distance measured at the front plane of the Image Receptor housing between the center of each pixel. Specified by a numeric pair - row spacing value (delimiter) column spacing value - in mm. | Table C.8-2. CR Image Module Attributes
Generator Power | (0018,1170) | 3 | Power in kW to the x-ray generator. | Table C.8-2. CR Image Module Attributes
Acquisition Device Processing Description | (0018,1400) | 3 | Describes device-specific processing associated with the image (e.g., Organ Description) | Table C.8-2. CR Image Module Attributes
Acquisition Device Processing Code | (0018,1401) | 3 | Code representing the device-specific processing associated with the image (e.g., CR Organ Filtering code) | Table C.8-2. CR Image Module Attributes
Cassette Orientation | (0018,1402) | 3 | Orientation of cassette, used to properly position the image for display. | Table C.8-2. CR Image Module Attributes
Cassette Size | (0018,1403) | 3 | Size of cassette. | Table C.8-2. CR Image Module Attributes
Exposures on Plate | (0018,1404) | 3 | Total number of x-ray exposures that have been made on the plate identified in Plate ID (0018,1004) | Table C.8-2. CR Image Module Attributes
Relative X-Ray Exposure | (0018,1405) | 3 | Relative x-ray exposure on the plate. Meaning of values is implementation specific. May be used to adjust the dynamic range of the plate digitizer (scanner). | Table C.8-2. CR Image Module Attributes
Sensitivity | (0018,6000) | 3 | Read out sensitivity. | Table C.8-2. CR Image Module Attributes
Pixel Spacing | (0028,0030) | 1C | Physical distance in the patient between the center of each pixel, specified by a numeric pair - adjacent row spacing (delimiter) adjacent column spacing in mm. | Table 10-10 Basic Pixel Spacing Calibration Macro Attributes
Pixel Spacing Calibration Type | (0028,0A02) | 3 | The type of correction for the effect of geometric magnification or calibration against an object of known size, if any. | Table 10-10 Basic Pixel Spacing Calibration Macro Attributes
Pixel Spacing Calibration Description | (0028,0A04) | 1C | A free text description of the type of correction or calibration performed. | Table 10-10 Basic Pixel Spacing Calibration Macro Attributes
Exposure Index | (0018,1411) | 3 | Measure of the detector response to radiation in the relevant image region of an image acquired with a digital x-ray imaging system as defined in IEC 62494-1. | Table 10-23 Exposure Index Macro Attributes
Target Exposure Index | (0018,1412) | 3 | The target value used to calculate Deviation Index (0018,1413) as defined in IEC 62494-1. | Table 10-23 Exposure Index Macro Attributes
Deviation Index | (0018,1413) | 3 | A scaled representation of the difference of the Exposure Index compared to the Target Exposure Index as defined in IEC 62494-1 and the report of AAPM TG 116. | Table 10-23 Exposure Index Macro Attributes
Anatomic Region Sequence | (0008,2218) | 3 | Sequence that identifies the anatomic region of interest in this Instance (i.e., external anatomy, surface anatomy, or general region of the body). | Table 10-7 General Anatomy Optional Macro Attributes
Anatomic Region Modifier Sequence | (0008,2220) | 3 | Sequence of Items that modifies the anatomic region of interest of this Instance. | Table 10-7 General Anatomy Optional Macro Attributes
Primary Anatomic Structure Sequence | (0008,2228) | 3 | Sequence of Items that identifies the primary anatomic structure(s) of interest in this Instance. | Table 10-8 Primary Anatomic Structure Macro Attributes
Primary Anatomic Structure Modifier Sequence | (0008,2230) | 3 | Sequence of Items that modifies the primary anatomic structure of interest in this Instance. | Table 10-8 Primary Anatomic Structure Macro Attributes