## Sheet1
IOD | IE | Module | Tag | Attribute Name | Keyword | Type | VR | VM | Attribute Description | Source | Standard Terms | Type_Group
Computed Radiography Image IOD | Image | CR_Image | (0028,0004) | Photometric Interpretation | Photometric​Interpretation | 1 | CS | 1 | Specifies the intended interpretation of the pixel data. | Table C.8-2. CR Image Module Attributes | {'Enumerated Values': ['MONOCHROME1', 'MONOCHROME2']} | Required
Computed Radiography Image IOD | Image | CR_Image | (0018,0060) | KVP | KVP | 3 | DS | 1 | Peak kilo voltage output of the X-Ray generator used. | Table C.8-2. CR Image Module Attributes |  | Optional
Computed Radiography Image IOD | Image | CR_Image | (0018,1004) | Plate ID | Plate​ID | 3 | LO | 1 | The ID or serial number of the sensing plate upon which the image was acquired. | Table C.8-2. CR Image Module Attributes |  | Optional
Computed Radiography Image IOD | Image | CR_Image | (0018,1110) | Distance Source to Detector | Distance​Source​To​Detector | 3 | DS | 1 | Distance in mm from source to detector center. | Table C.8-2. CR Image Module Attributes |  | Optional
Computed Radiography Image IOD | Image | CR_Image | (0018,1111) | Distance Source to Patient | Distance​Source​To​Patient | 3 | DS | 1 | Distance in mm from source to center of field of view. | Table C.8-2. CR Image Module Attributes |  | Optional
Computed Radiography Image IOD | Image | CR_Image | (0018,1150) | Exposure Time | Exposure​Time | 3 | IS | 1 | Time of X-Ray exposure in msec. | Table C.8-2. CR Image Module Attributes |  | Optional
Computed Radiography Image IOD | Image | CR_Image | (0018,1151) | X-Ray Tube Current | X​Ray​Tube​Current | 3 | IS | 1 | X-Ray Tube Current in mA. | Table C.8-2. CR Image Module Attributes |  | Optional
Computed Radiography Image IOD | Image | CR_Image | (0018,1152) | Exposure | Exposure | 3 | IS | 1 | The exposure expressed in mAs, for example calculated from Exposure Time and X-Ray Tube Current. | Table C.8-2. CR Image Module Attributes |  | Optional
Computed Radiography Image IOD | Image | CR_Image | (0018,1153) | Exposure in µAs | Exposure​Inu​As | 3 | IS | 1 | The exposure expressed in µAs, for example calculated from Exposure Time and X-Ray Tube Current. | Table C.8-2. CR Image Module Attributes |  | Optional
Computed Radiography Image IOD | Image | CR_Image | (0018,1164) | Imager Pixel Spacing | Imager​Pixel​Spacing | 3 | DS | 2 | Physical distance measured at the front plane of the Image Receptor housing between the center of each pixel. | Table C.8-2. CR Image Module Attributes |  | Optional
Computed Radiography Image IOD | Image | CR_Image | (0018,1170) | Generator Power | Generator​Power | 3 | IS | 1 | Power in kW to the X-Ray generator. | Table C.8-2. CR Image Module Attributes |  | Optional
Computed Radiography Image IOD | Image | CR_Image | (0018,1400) | Acquisition Device Processing Description | Acquisition​Device​Processing​Description | 3 | LO | 1 | Describes device-specific processing associated with the image (e.g., Organ Description). | Table C.8-2. CR Image Module Attributes |  | Optional
Computed Radiography Image IOD | Image | CR_Image | (0018,1401) | Acquisition Device Processing Code | Acquisition​Device​Processing​Code | 3 | LO | 1 | Code representing the device-specific processing associated with the image (e.g., CR Organ Filtering code). | Table C.8-2. CR Image Module Attributes |  | Optional
Computed Radiography Image IOD | Image | CR_Image | (0018,1402) | Cassette Orientation | Cassette​Orientation | 3 | CS | 1 | Orientation of cassette, used to properly position the image for display. | Table C.8-2. CR Image Module Attributes | {'Enumerated Values': ['LANDSCAPE', 'PORTRAIT']} | Optional
Computed Radiography Image IOD | Image | CR_Image | (0018,1403) | Cassette Size | Cassette​Size | 3 | CS | 1 | Size of cassette. | Table C.8-2. CR Image Module Attributes | {'Defined Terms': ['18CMX24CM', '8INX10IN', '24CMX30CM', '10INX12IN', '30CMX35CM', '30CMX40CM', '11INX14IN', '35CMX35CM', '14INX14IN', '35CMX43CM', '14INX17IN']} | Optional
Computed Radiography Image IOD | Image | CR_Image | (0018,1404) | Exposures on Plate | Exposures​On​Plate | 3 | US | 1 | Total number of X-Ray exposures that have been made on the plate identified in Plate ID (0018,1004). | Table C.8-2. CR Image Module Attributes |  | Optional
Computed Radiography Image IOD | Image | CR_Image | (0018,1405) | Relative X-Ray Exposure | Relative​X​Ray​Exposure | 3 | IS | 1 | Relative X-Ray exposure on the plate. Meaning of Values is implementation specific. | Table C.8-2. CR Image Module Attributes |  | Optional
Computed Radiography Image IOD | Image | CR_Image | (0018,6000) | Sensitivity | Sensitivity | 3 | DS | 1 | Read out sensitivity. | Table C.8-2. CR Image Module Attributes |  | Optional
Computed Radiography Image IOD | Image | CR_Image | (0028,0030) | Pixel Spacing | Pixel​Spacing | 1C | DS | 2 | Physical distance in the Patient between the center of each pixel, specified by a numeric pair - adjacent row spacing (delimiter) adjacent column spacing in mm. | Table 10-10 Basic Pixel Spacing Calibration Macro Attributes |  | Conditional
Computed Radiography Image IOD | Image | CR_Image | (0028,0A02) | Pixel Spacing Calibration Type | Pixel​Spacing​Calibration​Type | 3 | CS | 1 | The type of correction for the effect of geometric magnification or calibration against an object of known size, if any. | Table 10-10 Basic Pixel Spacing Calibration Macro Attributes | {'Enumerated Values': ['GEOMETRY', 'FIDUCIAL']} | Optional
Computed Radiography Image IOD | Image | CR_Image | (0028,0A04) | Pixel Spacing Calibration Description | Pixel​Spacing​Calibration​Description | 1C | LO | 1 | A free text description of the type of correction or calibration performed. | Table 10-10 Basic Pixel Spacing Calibration Macro Attributes |  | Conditional
Computed Radiography Image IOD | Image | CR_Image | (0008,2218) | Anatomic Region Sequence | Anatomic​Region​Sequence | 3 | SQ | 1 | Sequence that identifies the anatomic region of interest in this Instance (i.e., external anatomy, surface anatomy, or general region of the body). | Table 10-7 General Anatomy Optional Macro Attributes |  | Optional
Computed Radiography Image IOD | Image | CR_Image | (0008,2228) | Primary Anatomic Structure Sequence | Primary​Anatomic​Structure​Sequence | 3 | SQ | 1 | Sequence of Items that identifies the primary anatomic structure(s) of interest in this Instance. | Table 10-8 Primary Anatomic Structure Macro Attributes |  | Optional
Computed Radiography Image IOD | Image | CR_Image | (0018,1411) | Exposure Index | Exposure​Index | 3 | DS | 1 | Measure of the detector response to radiation in the relevant image region of an image acquired with a digital X-Ray imaging system as defined in IEC 62494-1. | Table 10-23 Exposure Index Macro Attributes |  | Optional
Computed Radiography Image IOD | Image | CR_Image | (0018,1412) | Target Exposure Index | Target​Exposure​Index | 3 | DS | 1 | The target value used to calculate Deviation Index (0018,1413) as defined in IEC 62494-1. | Table 10-23 Exposure Index Macro Attributes |  | Optional
Computed Radiography Image IOD | Image | CR_Image | (0018,1413) | Deviation Index | Deviation​Index | 3 | DS | 1 | A scaled representation of the difference of the Exposure Index compared to the Target Exposure Index as defined in IEC 62494-1 and the report of AAPM TG 116. | Table 10-23 Exposure Index Macro Attributes |  | Optional
CT Image IOD | Image | CT_Image | (0008,0008) | Image Type | Image​Type | 1 | CS | 2-n | Image identification characteristics. See Section C.8.2.1.1.1 for specialization. | Table C.8-3. CT Image Module Attributes | {'Enumerated Values' : ['ORIGINAL ', 'DERIVED', 'PRIMARY ', 'SECONDARY', 'AXIAL', 'LOCALIZER']} | Required
CT Image IOD | Image | CT_Image | (0018,9361) | Multi-energy CT Acquisition | Multienergy​CT​Acquisition | 3 | CS | 1 | Indicates whether the image is created by means of Multi-energy technique. | Table C.8-3. CT Image Module Attributes | {'Enumerated Values': ['YES', 'NO']} | Optional
CT Image IOD | Image | CT_Image | (0028,0002) | Samples per Pixel | Samples​Per​Pixel | 1 | US | 1 | Number of samples (planes) in this image. See Section C.8.2.1.1.2 for specialization. | Table C.8-3. CT Image Module Attributes |  | Required
CT Image IOD | Image | CT_Image | (0028,0004) | Photometric Interpretation | Photometric​Interpretation | 1 | CS | 1 | Specifies the intended interpretation of the pixel data. See Section C.8.2.1.1.3 for specialization. | Table C.8-3. CT Image Module Attributes | {'Enumerated Values': ['MONOCHROME1', 'MONOCHROME2']} | Required
CT Image IOD | Image | CT_Image | (0028,0100) | Bits Allocated | Bits​Allocated | 1 | US | 1 | Number of bits allocated for each pixel sample. Each sample shall have the same number of bits allocated. See Section C.8.2.1.1.4 for specialization. | Table C.8-3. CT Image Module Attributes |  | Required
CT Image IOD | Image | CT_Image | (0028,0101) | Bits Stored | Bits​Stored | 1 | US | 1 | Number of bits stored for each pixel sample. Each sample shall have the same number of bits stored. See Section C.8.2.1.1.5 for specialization. | Table C.8-3. CT Image Module Attributes |  | Required
CT Image IOD | Image | CT_Image | (0028,0102) | High Bit | High​Bit | 1 | US | 1 | Most significant bit for pixel sample data. Each sample shall have the same high bit. See Section C.8.2.1.1.6 for specialization. | Table C.8-3. CT Image Module Attributes |  | Required
CT Image IOD | Image | CT_Image | (0028,1052) | Rescale Intercept | Rescale​Intercept | 1 | DS | 1 | The value b in relationship between stored values (SV) and the output units. | Table C.8-3. CT Image Module Attributes |  | Required
CT Image IOD | Image | CT_Image | (0028,1053) | Rescale Slope | Rescale​Slope | 1 | DS | 1 | m in the equation specified in Rescale Intercept (0028,1052). | Table C.8-3. CT Image Module Attributes |  | Required
CT Image IOD | Image | CT_Image | (0028,1054) | Rescale Type | Rescale​Type | 1C | LO | 1 | Specifies the output units of Rescale Slope (0028,1053) and Rescale Intercept (0028,1052). | Table C.8-3. CT Image Module Attributes |  | Conditional
CT Image IOD | Image | CT_Image | (0018,0060) | KVP | KVP | 2 | DS | 1 | Peak kilo voltage output of the X-Ray generator used. | Table C.8-3. CT Image Module Attributes |  | Required
CT Image IOD | Image | CT_Image | (0020,0012) | Acquisition Number | Acquisition​Number | 2 | IS | 1 | A number identifying the single continuous gathering of data over a period of time that resulted in this image. | Table C.8-3. CT Image Module Attributes |  | Required
CT Image IOD | Image | CT_Image | (0018,0022) | Scan Options | Scan​Options | 3 | CS_undefined | 1-n | Parameters of scanning sequence. | Table C.8-3. CT Image Module Attributes |  | Optional
CT Image IOD | Image | CT_Image | (0018,0090) | Data Collection Diameter | Data​Collection​Diameter | 3 | DS | 1 | The diameter in mm of the region over which data were collected. | Table C.8-3. CT Image Module Attributes |  | Optional
CT Image IOD | Image | CT_Image | (0018,9313) | Data Collection Center (Patient) | Data​Collection​Center​Patient | 3 | FD | 3 | The x, y, and z coordinates (in the Patient-Based Coordinate System) in mm of the center of the region in which data were collected. See Section C.8.15.3.6.1. | Table C.8-3. CT Image Module Attributes |  | Optional
CT Image IOD | Image | CT_Image | (0018,1100) | Reconstruction Diameter | Reconstruction​Diameter | 3 | DS | 1 | Diameter in mm of the region from within which data were used in creating the reconstruction of the image. | Table C.8-3. CT Image Module Attributes |  | Optional
CT Image IOD | Image | CT_Image | (0018,9318) | Reconstruction Target Center (Patient) | Reconstruction​Target​Center​Patient | 3 | FD | 3 | The x, y, and z coordinates (in the Patient-Based Coordinate System) of the reconstruction center target point as used for reconstruction in mm. | Table C.8-3. CT Image Module Attributes |  | Optional
CT Image IOD | Image | CT_Image | (0018,1110) | Distance Source to Detector | Distance​Source​To​Detector | 3 | DS | 1 | Distance in mm from source to detector center. | Table C.8-3. CT Image Module Attributes |  | Optional
CT Image IOD | Image | CT_Image | (0018,1111) | Distance Source to Patient | Distance​Source​To​Patient | 3 | DS | 1 | Distance in mm from source to isocenter (center of field of view). | Table C.8-3. CT Image Module Attributes |  | Optional
CT Image IOD | Image | CT_Image | (0018,1120) | Gantry/Detector Tilt | Gantry​Detector​Tilt | 3 | DS | 1 | Nominal angle of tilt in degrees of the scanning gantry. Not intended for mathematical computations. | Table C.8-3. CT Image Module Attributes |  | Optional
CT Image IOD | Image | CT_Image | (0018,1130) | Table Height | Table​Height | 3 | DS | 1 | The distance in mm of the top of the patient table to the center of rotation; below the center is positive. | Table C.8-3. CT Image Module Attributes |  | Optional
CT Image IOD | Image | CT_Image | (0018,1140) | Rotation Direction | Rotation​Direction | 3 | CS | 1 | Direction of rotation of the source when relevant, about nearest principal axis of equipment. | Table C.8-3. CT Image Module Attributes | {'Enumerated Values': ['CW', 'CC']} | Optional
CT Image IOD | Image | CT_Image | (0018,1150) | Exposure Time | Exposure​Time | 3 | IS | 1 | Time of X-Ray exposure in msec. | Table C.8-3. CT Image Module Attributes |  | Optional
CT Image IOD | Image | CT_Image | (0018,1151) | X-Ray Tube Current | X​Ray​Tube​Current | 3 | IS | 1 | X-Ray Tube Current in mA. | Table C.8-3. CT Image Module Attributes |  | Optional
CT Image IOD | Image | CT_Image | (0018,1152) | Exposure | Exposure | 3 | IS | 1 | The exposure expressed in mAs, for example calculated from Exposure Time and X-Ray Tube Current. | Table C.8-3. CT Image Module Attributes |  | Optional
CT Image IOD | Image | CT_Image | (0018,1153) | Exposure in µAs | Exposure​Inu​As | 3 | IS | 1 | The exposure expressed in µAs, for example calculated from Exposure Time and X-Ray Tube Current. | Table C.8-3. CT Image Module Attributes |  | Optional
CT Image IOD | Image | CT_Image | (0018,1160) | Filter Type | Filter​Type | 3 | SH | 1 | Label for the type of filter inserted into the X-Ray beam. | Table C.8-3. CT Image Module Attributes |  | Optional
CT Image IOD | Image | CT_Image | (0018,7050) | Filter Material | Filter​Material | 3 | CS | 1-n | The X-Ray absorbing material used in the filter. May be multi-valued. | Table C.8-3. CT Image Module Attributes | {'Defined Terms': ['MOLYBDENUM', 'ALUMINUM', 'COPPER', 'RHODIUM', 'NIOBIUM', 'EUROPIUM', 'LEAD', 'TANTALUM', 'SILVER', 'TIN']} | Optional
CT Image IOD | Image | CT_Image | (0018,1170) | Generator Power | Generator​Power | 3 | IS | 1 | Power in kW to the X-Ray generator. | Table C.8-3. CT Image Module Attributes |  | Optional
CT Image IOD | Image | CT_Image | (0018,1190) | Focal Spot(s) | Focal​Spots | 3 | DS | 1-n | Size of the focal spot in mm. For devices with variable focal spot or multiple focal spots, small dimension followed by large dimension. | Table C.8-3. CT Image Module Attributes |  | Optional
CT Image IOD | Image | CT_Image | (0018,1210) | Convolution Kernel | Convolution​Kernel | 3 | SH | 1-n | A label describing the convolution kernel or algorithm used to reconstruct the data. | Table C.8-3. CT Image Module Attributes |  | Optional
CT Image IOD | Image | CT_Image | (0018,9305) | Revolution Time | Revolution​Time | 3 | FD | 1 | The time in seconds of a complete revolution of the source around the gantry orbit. | Table C.8-3. CT Image Module Attributes |  | Optional
CT Image IOD | Image | CT_Image | (0018,9306) | Single Collimation Width | Single​Collimation​Width | 3 | FD | 1 | The width of a single row of acquired data (in mm). | Table C.8-3. CT Image Module Attributes |  | Optional
CT Image IOD | Image | CT_Image | (0018,9307) | Total Collimation Width | Total​Collimation​Width | 3 | FD | 1 | The width of the total collimation (in mm) over the area of active X-Ray detection. | Table C.8-3. CT Image Module Attributes |  | Optional
CT Image IOD | Image | CT_Image | (0018,9309) | Table Speed | Table​Speed | 3 | FD | 1 | The distance in mm that the table moves in one second during the gathering of data that resulted in this image. | Table C.8-3. CT Image Module Attributes |  | Optional
CT Image IOD | Image | CT_Image | (0018,9310) | Table Feed per Rotation | Table​Feed​Per​Rotation | 3 | FD | 1 | Motion of the table (in mm) during a complete revolution of the source around the gantry orbit. | Table C.8-3. CT Image Module Attributes |  | Optional
CT Image IOD | Image | CT_Image | (0018,9311) | Spiral Pitch Factor | Spiral​Pitch​Factor | 3 | FD | 1 | Ratio of the Table Feed per Rotation (0018,9310) to the Total Collimation Width (0018,9307). | Table C.8-3. CT Image Module Attributes |  | Optional
CT Image IOD | Image | CT_Image | (0018,9323) | Exposure Modulation Type | Exposure​Modulation​Type | 3 | CS | 1-n | A label describing the type of exposure modulation used for the purpose of limiting the dose. | Table C.8-3. CT Image Module Attributes | {'Defined Terms': ['NONE']} | Optional
CT Image IOD | Image | CT_Image | (0018,9345) | CTDIvol | CTD​Ivol | 3 | FD | 1 | Computed Tomography Dose Index (CTDIvol), in mGy according to the principles described in [IEC 60601-2-44]. | Table C.8-3. CT Image Module Attributes |  | Optional
CT Image IOD | Image | CT_Image | (0018,9346) | CTDI Phantom Type Code Sequence | CTDI​Phantom​Type​Code​Sequence | 3 | SQ | 1 | The type of phantom used for CTDI measurement. | Table C.8-3. CT Image Module Attributes |  | Optional
CT Image IOD | Image | CT_Image | (0018,1271) | Water Equivalent Diameter | Water​Equivalent​Diameter | 3 | FD | 1 | The diameter, in mm, of a cylinder of water having the same X-Ray attenuation as the patient for this reconstructed slice. | Table C.8-3. CT Image Module Attributes |  | Optional
CT Image IOD | Image | CT_Image | (0018,1272) | Water Equivalent Diameter Calculation Method Code Sequence | Water​Equivalent​Diameter​Calculation​Method​Code​Sequence | 1C | SQ | 1 | The method of calculation of Water Equivalent Diameter (0018,1271). | Table C.8-3. CT Image Module Attributes |  | Conditional
CT Image IOD | Image | CT_Image | (0018,115E) | Image and Fluoroscopy Area Dose Product | Image​And​Fluoroscopy​Area​Dose​Product | 3 | DS | 1 | X-Ray dose, measured in dGy*cm*cm, to which the patient was exposed for the acquisition of the entire Irradiation Event from which this image was reconstructed. | Table C.8-3. CT Image Module Attributes |  | Optional
CT Image IOD | Image | CT_Image | (0018,9351) | Calcium Scoring Mass Factor Patient | Calcium​Scoring​Mass​Factor​Patient | 3 | FL | 1 | The calibration factor for the calcium mass score. | Table C.8-3. CT Image Module Attributes |  | Optional
CT Image IOD | Image | CT_Image | (0018,9352) | Calcium Scoring Mass Factor Device | Calcium​Scoring​Mass​Factor​Device | 3 | FL | 3 | The calibration factors for the calcium mass score of the device. | Table C.8-3. CT Image Module Attributes |  | Optional
CT Image IOD | Image | CT_Image | (0018,9353) | Energy Weighting Factor | Energy​Weighting​Factor | 1C | FL | 1 | The weighting factor of the data from the primary source in a multiple energy composition image. | Table C.8-3. CT Image Module Attributes |  | Conditional
CT Image IOD | Image | CT_Image | (0018,9360) | CT Additional X-Ray Source Sequence | CT​Additional​X​Ray​Source​Sequence | 3 | SQ | 1 | Acquisition parameters for X-Ray sources beyond the primary X-Ray source, which is specified in other Attributes of this Module. | Table C.8-3. CT Image Module Attributes |  | Optional
CT Image IOD | Image | CT_Image | (0008,2218) | Anatomic Region Sequence | Anatomic​Region​Sequence | 3 | SQ | 1 | Sequence that identifies the anatomic region of interest in this Instance (i.e., external anatomy, surface anatomy, or general region of the body). Only a single Item is permitted in this Sequence. | Table 10-7. General Anatomy Optional Macro Attributes |  | Optional
CT Image IOD | Image | CT_Image | (0008,2228) | Primary Anatomic Structure Sequence | Primary​Anatomic​Structure​Sequence | 3 | SQ | 1 | Sequence of Items that identifies the primary anatomic structure(s) of interest in this Instance. | Table 10-8 Primary Anatomic Structure Macro Attributes |  | Optional
CT Image IOD | Image | CT_Image | (0054,0220) | View Code Sequence | View​Code​Sequence | 3 | SQ | 1 | Sequence that describes the projection of the anatomic region of interest. Only a single Item is permitted in this Sequence. | Table 10-25. Optional View and Slice Progression Direction Macro Attributes |  | Optional
CT Image IOD | Image | CT_Image | (0054,0500) | Slice Progression Direction | Slice​Progression​Direction | 3 | CS | 1 | Describes the anatomical direction in which a set of slices is progressing (see Section 10.20.1.1). Meaningful only for cardiac images. Enumerated Values are defined in Section 10.20.1.1. | Table 10-25. Optional View and Slice Progression Direction Macro Attributes | {'Enumerated Values': ['APEX_TO_BASE', 'BASE_TO_APEX', 'ANT_TO_INF', 'INF_TO_ANT', 'SEPTUM_TO_WALL', 'WALL_TO_SEPTUM']} | Optional
CT Image IOD | Image | CT_Image | (300A,0675) | Equipment Frame of Reference UID | Equipment​Frame​Of​Reference​UID | 1C | UI | 1 | Frame of Reference identifier for the Treatment Delivery Device. Required if either Imaging Equipment to Treatment Delivery Device Relationship Sequence (300A,07A1) or Patient to Equipment Relationship Sequence (300A,07A0) is present. See Section C.36.12.1 and Section C.36.2.4.12.1.1. | Table C.36.2.4.12-1 RT Equipment Mapping and Plan Reference Macro Attributes |  | Conditional
CT Image IOD | Image | CT_Image | (300A,07A1) | Imaging Equipment to Treatment Delivery Device Relationship Sequence | Imaging​Equipment​To​Treatment​Delivery​Device​Relationship​Sequence | 3 | SQ | 1 | The relationship between the coordinate system of the Imaging Device and coordinate system of the Treatment Delivery Device. Only a single Item is permitted in this Sequence. See Section 10.39.1.1. | Table C.36.2.4.12-1 RT Equipment Mapping and Plan Reference Macro Attributes |  | Optional
CT Image IOD | Image | CT_Image | (300A,012C) | Isocenter Position | Isocenter​Position | 3 | DS | 3 | Isocenter coordinates (x,y,z), in mm. Specifies the location of the machine isocenter in the Patient-Based Coordinate System associated with the Frame of Reference. It allows transformation from the Equipment-Based Coordinate System to the Patient-Based Coordinate System. | Table C.36.2.4.12-1 RT Equipment Mapping and Plan Reference Macro Attributes |  | Optional
CT Image IOD | Image | CT_Image | (300A,07A0) | Patient to Equipment Relationship Sequence | Patient​To​Equipment​Relationship​Sequence | 3 | SQ | 1 | The relationship between the Patient Frame of Reference and the coordinate system of the Treatment Delivery Device. Only a single Item is permitted in this Sequence. | Table C.36.2.4.12-1 RT Equipment Mapping and Plan Reference Macro Attributes |  | Optional
CT Image IOD | Image | CT_Image | (300A,079F) | Patient Treatment Preparation Sequence | Patient​Treatment​Preparation​Sequence | 3 | SQ | 1 | Procedures and devices required to prepare a patient for treatment. Only a single Item is permitted in this Sequence. See Section C.8.8.12.1.3. | Table C.36.2.4.12-1 RT Equipment Mapping and Plan Reference Macro Attributes |  | Optional
CT Image IOD | Image | CT_Image | (300C,0002) | Referenced RT Plan Sequence | Referenced​RT​Plan​Sequence | 3 | SQ | 1 | RT Plan Instance associated with the acquisition of this Instance. Only a single Item is permitted in this Sequence. | Table C.36.2.4.12-1 RT Equipment Mapping and Plan Reference Macro Attributes |  | Optional
CT Image IOD | Image | CT_Image | (3002,012E) | Scan Arc Type | Scan​Arc​Type | 3 | CS | 1 | Categorization of the amount of rotation of the scan. | Table C.36.2.4.5-2 3D RT Cone-Beam Imaging Geometry Optional Macro Attributes | {'Enumerated Values': ['FULL_ARC', 'HALF_ARC', 'CUSTOM_ARC']} | Optional
CT Image IOD | Image | CT_Image | (3002,012B) | Scan Start Position Sequence | Scan​Start​Position​Sequence | 3 | SQ | 1 | Start Position of the 3D RT Cone Beam Radiation Imaging Acquisition. The angle is a Continuous Rotation Angle; see Section C.36.1.1.5. | Table C.36.2.4.5-2 3D RT Cone-Beam Imaging Geometry Optional Macro Attributes |  | Optional
CT Image IOD | Image | CT_Image | (3002,012C) | Scan Stop Position Sequence | Scan​Stop​Position​Sequence | 3 | SQ | 1 | Stop Position of the 3D RT Cone Beam Radiation Imaging Acquisition. The angle is a Continuous Rotation Angle; see Section C.36.1.1.5. | Table C.36.2.4.5-2 3D RT Cone-Beam Imaging Geometry Optional Macro Attributes |  | Optional
CT Image IOD | Image | CT_Image | (3002,012F) | Detector Positioning Type | Detector​Positioning​Type | 3 | CS | 1 | Fan type of acquisition. Enumerated Values: CENTERED: Full fan, detector is centered, resulting in a normal field of view. SHIFTED: Half fan, detector is laterally shifted by half of the detector size to increase the field of view. | Table C.36.2.4.5-2 3D RT Cone-Beam Imaging Geometry Optional Macro Attributes | {'Enumerated Values': ['CENTERED', 'SHIFTED']} | Optional
Digital Mammography X-Ray Image IOD | Image | DX_Anatomy_Imaged | (0020,0062) | Image Laterality | Image​Laterality | 1 | CS | 1 | Laterality of (possibly paired) body part (as described in Anatomic Region Sequence (0008,2218)) examined. | Table C.8-69. DX Anatomy Imaged Module Attributes | {'Enumerated Values': ['R', 'L', 'U', 'B']} | Required
Digital Mammography X-Ray Image IOD | Image | DX_Anatomy_Imaged | (0008,2218) | Anatomic Region Sequence | Anatomic​Region​Sequence | 2 | SQ | 1 | Sequence that identifies the anatomic region of interest in this Instance (i.e., external anatomy, surface anatomy, or general region of the body). Zero or one Item shall be included in this Sequence. | Table 10-6. General Anatomy Required Macro Attributes |  | Required
Digital Mammography X-Ray Image IOD | Image | DX_Anatomy_Imaged | (0008,2220) | Anatomic Region Modifier Sequence | Anatomic​Region​Modifier​Sequence | 3 | SQ | 1 | Sequence of Items that modifies the anatomic region of interest of this Instance. One or more Items are permitted in this Sequence. | Table 10-6. General Anatomy Required Macro Attributes |  | Optional
Digital Mammography X-Ray Image IOD | Image | DX_Anatomy_Imaged | (0008,2228) | Primary Anatomic Structure Sequence | Primary​Anatomic​Structure​Sequence | 3 | SQ | 1 | Sequence of Items that identifies the primary anatomic structure(s) of interest in this Instance. One or more Items are permitted in this Sequence. | Table 10-8. Primary Anatomic Structure Macro Attributes |  | Optional
Digital Mammography X-Ray Image IOD | Image | DX_Detector | (0018,7014) | Detector Active Time | Detector​Active​Time | 3 | DS | 1 | Time in mSec that the detector is active during acquisition of this image. | Table C.8-71 DX Detector Module Attributes |  | Optional
Digital Mammography X-Ray Image IOD | Image | DX_Detector | (0018,7016) | Detector Activation Offset From Exposure | Detector​Activation​Offset​From​Exposure | 3 | DS | 1 | Offset time in mSec that the detector becomes active after the X-Ray beam is turned on during acquisition of this image. May be negative. | Table C.8-71 DX Detector Module Attributes |  | Optional
Digital Mammography X-Ray Image IOD | Image | DX_Detector | (0018,1147) | Field of View Shape | Field​Of​View​Shape | 3 | CS | 1 | Shape of the Field of View, that is the image pixels stored in Pixel Data (7FE0,0010). | Table C.8-71 DX Detector Module Attributes | {'Enumerated Values': ['RECTANGLE', 'ROUND', 'HEXAGONAL']} | Optional
Digital Mammography X-Ray Image IOD | Image | DX_Detector | (0018,1149) | Field of View Dimension(s) | Field​Of​View​Dimensions | 3 | IS | 1-2 | Dimensions in mm of the Field of View, that is the image pixels stored in Pixel Data (7FE0,0010). | Table C.8-71 DX Detector Module Attributes |  | Optional
Digital Mammography X-Ray Image IOD | Image | DX_Detector | (0018,7030) | Field of View Origin | Field​Of​View​Origin | 1C | DS | 2 | Offset of the TLHC of a rectangle circumscribing the Field of View, that is the image pixels stored in Pixel Data (7FE0,0010), before rotation or flipping, from the TLHC of the physical detector area measured in physical detector pixels as a row offset followed by a column offset. | Table C.8-71 DX Detector Module Attributes |  | Conditional
Digital Mammography X-Ray Image IOD | Image | DX_Detector | (0018,7032) | Field of View Rotation | Field​Of​View​Rotation | 1C | DS | 1 | Clockwise rotation in degrees of Field of View, that is the image pixels stored in Pixel Data (7FE0,0010), relative to the physical detector. | Table C.8-71 DX Detector Module Attributes |  | Conditional
Digital Mammography X-Ray Image IOD | Image | DX_Detector | (0018,7034) | Field of View Horizontal Flip | Field​Of​View​Horizontal​Flip | 1C | CS | 1 | Whether or not a horizontal flip has been applied to the Field of View, that is the image pixels stored in Pixel Data (7FE0,0010), after rotation relative to the physical detector as described in Field of View Rotation (0018,7032). | Table C.8-71 DX Detector Module Attributes | {'Enumerated Values': ['YES', 'NO']} | Conditional
Digital Mammography X-Ray Image IOD | Image | DX_Detector | (0018,1164) | Imager Pixel Spacing | Imager​Pixel​Spacing | 1 | DS | 2 | Physical distance measured at the front plane of the detector housing between the center of each image pixel specified by a numeric pair - row spacing Value(delimiter) column spacing Value in mm. | Table C.8-71 DX Detector Module Attributes |  | Required
Digital Mammography X-Ray Image IOD | Image | DX_Detector | (0028,0030) | Pixel Spacing | Pixel​Spacing | 1C | DS | 2 | Physical distance in the Patient between the center of each pixel, specified by a numeric pair - adjacent row spacing (delimiter) adjacent column spacing in mm. See Section 10.7.1.1 and Section 10.7.1.3. Required if the image has been calibrated. May be present otherwise. | Table 10-10. Basic Pixel Spacing Calibration Macro Attributes |  | Conditional
Digital Mammography X-Ray Image IOD | Image | DX_Detector | (0028,0A02) | Pixel Spacing Calibration Type | Pixel​Spacing​Calibration​Type | 3 | CS | 1 | The type of correction for the effect of geometric magnification or calibration against an object of known size, if any. See Section 10.7.1.2. | Table 10-10. Basic Pixel Spacing Calibration Macro Attributes | {'Enumerated Values': ['GEOMETRY', 'FIDUCIAL']} | Optional
Digital Mammography X-Ray Image IOD | Image | DX_Detector | (0028,0A04) | Pixel Spacing Calibration Description | Pixel​Spacing​Calibration​Description | 1C | LO | 1 | A free text description of the type of correction or calibration performed. | Table 10-10. Basic Pixel Spacing Calibration Macro Attributes |  | Conditional
Digital Mammography X-Ray Image IOD | Image | DX_Detector | (0018,1007) | Cassette ID | Cassette​ID | 3 | LO | 1 | Identifier of the cassette that contains the photostimulable phosphor plate, for CR acquisitions. | Table C.8-71 DX Detector Module Attributes |  | Optional
Digital Mammography X-Ray Image IOD | Image | DX_Detector | (0018,1004) | Plate ID | Plate​ID | 3 | LO | 1 | Identifier of the photostimulable phosphor plate, for CR acquisitions. | Table C.8-71 DX Detector Module Attributes |  | Optional
Digital Mammography X-Ray Image IOD | Image | DX_Detector | (0018,7004) | Detector Type | Detector​Type | 2 | CS | 1 | The type of detector used to acquire this image. | Table C.8-71b Digital X-Ray Detector Macro Attributes | {'Defined Terms': ['DIRECT', 'SCINTILLATOR', 'STORAGE', 'FILM']} | Required
Digital Mammography X-Ray Image IOD | Image | DX_Detector | (0018,7005) | Detector Configuration | Detector​Configuration | 3 | CS | 1 | The physical configuration of the detector. | Table C.8-71b Digital X-Ray Detector Macro Attributes | {'Defined Terms': ['AREA', 'SLOT']} | Optional
Digital Mammography X-Ray Image IOD | Image | DX_Detector | (0018,7006) | Detector Description | Detector​Description | 3 | LT | 1 | Free text description of detector. | Table C.8-71b Digital X-Ray Detector Macro Attributes |  | Optional
Digital Mammography X-Ray Image IOD | Image | DX_Detector | (0018,7008) | Detector Mode | Detector​Mode | 3 | LT | 1 | Text description of operating mode of detector (implementation specific). | Table C.8-71b Digital X-Ray Detector Macro Attributes |  | Optional
Digital Mammography X-Ray Image IOD | Image | DX_Detector | (0018,700A) | Detector ID | Detector​ID | 3 | SH | 1 | The ID or serial number of the detector used to acquire this image. | Table C.8-71b Digital X-Ray Detector Macro Attributes |  | Optional
Digital Mammography X-Ray Image IOD | Image | DX_Detector | (0018,700C) | Date of Last Detector Calibration | Date​Of​Last​Detector​Calibration | 3 | DA | 1 | The date on which the detector used to acquire this image as identified in Detector ID (0018,700A) was last calibrated. | Table C.8-71b Digital X-Ray Detector Macro Attributes |  | Optional
Digital Mammography X-Ray Image IOD | Image | DX_Detector | (0018,700E) | Time of Last Detector Calibration | Time​Of​Last​Detector​Calibration | 3 | TM | 1 | The time at which the detector used to acquire this image as identified in Detector ID (0018,700A) was last calibrated. | Table C.8-71b Digital X-Ray Detector Macro Attributes |  | Optional
Digital Mammography X-Ray Image IOD | Image | DX_Detector | (0018,7010) | Exposures on Detector Since Last Calibration | Exposures​On​Detector​Since​Last​Calibration | 3 | IS | 1 | Total number of X-Ray exposures that have been made on the detector used to acquire this image as identified in Detector ID (0018,700A) since it was calibrated. | Table C.8-71b Digital X-Ray Detector Macro Attributes |  | Optional
Digital Mammography X-Ray Image IOD | Image | DX_Detector | (0018,7011) | Exposures on Detector Since Manufactured | Exposures​On​Detector​Since​Manufactured | 3 | IS | 1 | Total number of X-Ray exposures that have been made on the detector used to acquire this image as identified in Detector ID (0018,700A) since it was manufactured. | Table C.8-71b Digital X-Ray Detector Macro Attributes |  | Optional
Digital Mammography X-Ray Image IOD | Image | DX_Detector | (0018,7012) | Detector Time Since Last Exposure | Detector​Time​Since​Last​Exposure | 3 | DS | 1 | Time in Seconds since an exposure was last made on this detector prior to the acquisition of this image. | Table C.8-71b Digital X-Ray Detector Macro Attributes |  | Optional
Digital Mammography X-Ray Image IOD | Image | DX_Detector | (0018,701A) | Detector Binning | Detector​Binning | 3 | DS | 2 | Number of active detectors used to generate a single pixel. Specified as number of row detectors per pixel then column. | Table C.8-71b Digital X-Ray Detector Macro Attributes |  | Optional
Digital Mammography X-Ray Image IOD | Image | DX_Detector | (0018,702A) | Detector Manufacturer Name | Detector​Manufacturer​Name | 3 | LO | 1 | Name of the manufacturer of the detector component of the acquisition system. | Table C.8-71b Digital X-Ray Detector Macro Attributes |  | Optional
Digital Mammography X-Ray Image IOD | Image | DX_Detector | (0018,702B) | Detector Manufacturer's Model Name | Detector​Manufacturer​Model​Name | 3 | LO | 1 | Model name of the detector component of the acquisition system. | Table C.8-71b Digital X-Ray Detector Macro Attributes |  | Optional
Digital Mammography X-Ray Image IOD | Image | DX_Detector | (0018,7000) | Detector Conditions Nominal Flag | Detector​Conditions​Nominal​Flag | 3 | CS | 1 | Whether or not the detector is operating within normal tolerances during this image acquisition. | Table C.8-71b Digital X-Ray Detector Macro Attributes | {'Enumerated Values': ['YES', 'NO']} | Optional
Digital Mammography X-Ray Image IOD | Image | DX_Detector | (0018,7001) | Detector Temperature | Detector​Temperature | 3 | DS | 1 | Detector temperature during exposure in degrees Celsius. | Table C.8-71b Digital X-Ray Detector Macro Attributes |  | Optional
Digital Mammography X-Ray Image IOD | Image | DX_Detector | (0018,6000) | Sensitivity | Sensitivity | 3 | DS | 1 | Detector sensitivity in manufacturer specific units. | Table C.8-71b Digital X-Ray Detector Macro Attributes |  | Optional
Digital Mammography X-Ray Image IOD | Image | DX_Detector | (0018,7020) | Detector Element Physical Size | Detector​Element​Physical​Size | 3 | DS | 2 | Physical dimensions of each detector element that comprises the detector matrix, in mm. | Table C.8-71b Digital X-Ray Detector Macro Attributes |  | Optional
Digital Mammography X-Ray Image IOD | Image | DX_Detector | (0018,7022) | Detector Element Spacing | Detector​Element​Spacing | 3 | DS | 2 | Physical distance between the center of each detector element, specified by a numeric pair - row spacing Value(delimiter) column spacing Value in mm. | Table C.8-71b Digital X-Ray Detector Macro Attributes |  | Optional
Digital Mammography X-Ray Image IOD | Image | DX_Detector | (0018,7024) | Detector Active Shape | Detector​Active​Shape | 3 | CS | 1 | Shape of the active area. | Table C.8-71b Digital X-Ray Detector Macro Attributes | {'Enumerated Values': ['RECTANGLE', 'ROUND', 'HEXAGONAL']} | Optional
Digital Mammography X-Ray Image IOD | Image | DX_Detector | (0018,7026) | Detector Active Dimension(s) | Detector​Active​Dimensions | 3 | DS | 1-2 | Dimensions in mm of the active area. | Table C.8-71b Digital X-Ray Detector Macro Attributes |  | Optional
Digital Mammography X-Ray Image IOD | Image | DX_Detector | (0018,7028) | Detector Active Origin | Detector​Active​Origin | 3 | DS | 2 | Offset of the TLHC of a rectangle circumscribing the active detector area from the TLHC of a rectangle circumscribing the physical detector area, measured in physical detector pixels as a row offset followed by a column offset. | Table C.8-71b Digital X-Ray Detector Macro Attributes |  | Optional
Digital Mammography X-Ray Image IOD | Image | DX_Detector | (0018,1411) | Exposure Index | Exposure​Index | 3 | DS | 1 | Measure of the detector response to radiation in the relevant image region of an image acquired with a digital X-Ray imaging system as defined in IEC 62494-1. | Table 10-23 Exposure Index Macro Attributes |  | Optional
Digital Mammography X-Ray Image IOD | Image | DX_Detector | (0018,1412) | Target Exposure Index | Target​Exposure​Index | 3 | DS | 1 | The target value used to calculate Deviation Index (0018,1413) as defined in IEC 62494-1. | Table 10-23 Exposure Index Macro Attributes |  | Optional
Digital Mammography X-Ray Image IOD | Image | DX_Detector | (0018,1413) | Deviation Index | Deviation​Index | 3 | DS | 1 | A scaled representation of the difference of the Exposure Index compared to the Target Exposure Index as defined in IEC 62494-1 and the report of AAPM TG 116. | Table 10-23 Exposure Index Macro Attributes |  | Optional
Digital Mammography X-Ray Image IOD | Image | DX_Image | (0008,0008) | Image Type | Image​Type | 1 | CS | 2-n | Image identification characteristics. See Section C.8.11.3.1.1 for specialization. | Table C.8-70. DX Image Module Attributes | {'Enumerated Values' : ['ORIGINAL ', 'DERIVED', 'PRIMARY ', 'SECONDARY', '']} | Required
Digital Mammography X-Ray Image IOD | Image | DX_Image | (0028,0002) | Samples per Pixel | Samples​Per​Pixel | 1 | US | 1 | Number of samples in this image. Enumerated Values: 1 | Table C.8-70. DX Image Module Attributes |  | Required
Digital Mammography X-Ray Image IOD | Image | DX_Image | (0028,0004) | Photometric Interpretation | Photometric​Interpretation | 1 | CS | 1 | Specifies the intended interpretation of the pixel data. Enumerated Values: MONOCHROME1, MONOCHROME2 | Table C.8-70. DX Image Module Attributes | {'Enumerated Values': ['MONOCHROME1', 'MONOCHROME2']} | Required
Digital Mammography X-Ray Image IOD | Image | DX_Image | (0028,0100) | Bits Allocated | Bits​Allocated | 1 | US | 1 | Number of bits allocated for each pixel sample. Enumerated Values: 8, 16 | Table C.8-70. DX Image Module Attributes |  | Required
Digital Mammography X-Ray Image IOD | Image | DX_Image | (0028,0101) | Bits Stored | Bits​Stored | 1 | US | 1 | Number of bits stored for each pixel sample. Enumerated Values: 6, 7, 8, 9, 10, 11, 12, 13, 14, 15, 16 | Table C.8-70. DX Image Module Attributes |  | Required
Digital Mammography X-Ray Image IOD | Image | DX_Image | (0028,0102) | High Bit | High​Bit | 1 | US | 1 | Most significant bit for pixel sample data. Shall be one less than the Value in Bit Stored (0028,0101). | Table C.8-70. DX Image Module Attributes |  | Required
Digital Mammography X-Ray Image IOD | Image | DX_Image | (0028,0103) | Pixel Representation | Pixel​Representation | 1 | US | 1 | Data representation of the pixel samples. Enumerated Values: 0000H Unsigned Integer | Table C.8-70. DX Image Module Attributes |  | Required
Digital Mammography X-Ray Image IOD | Image | DX_Image | (0028,1040) | Pixel Intensity Relationship | Pixel​Intensity​Relationship | 1 | CS | 1 | The relationship between the pixel sample values and the X-Ray beam intensity. Enumerated Values: LIN, LOG | Table C.8-70. DX Image Module Attributes | {'Enumerated Values': ['LIN', 'LOG']} | Required
Digital Mammography X-Ray Image IOD | Image | DX_Image | (0028,1041) | Pixel Intensity Relationship Sign | Pixel​Intensity​Relationship​Sign | 1 | SS | 1 | The sign of the relationship between the pixel sample values stored in Pixel Data (7FE0,0010) and the X-Ray beam intensity. Enumerated Values: +1, -1 | Table C.8-70. DX Image Module Attributes |  | Required
Digital Mammography X-Ray Image IOD | Image | DX_Image | (0028,1052) | Rescale Intercept | Rescale​Intercept | 1 | DS | 1 | The value b in the relationship between stored values (SV) in Pixel Data (7FE0,0010) and the output units specified in Rescale Type (0028,1054). Output units = m*SV + b. Enumerated Values: 0 | Table C.8-70. DX Image Module Attributes |  | Required
Digital Mammography X-Ray Image IOD | Image | DX_Image | (0028,1053) | Rescale Slope | Rescale​Slope | 1 | DS | 1 | m in the equation specified by Rescale Intercept (0028,1052). Enumerated Values: 1 | Table C.8-70. DX Image Module Attributes |  | Required
Digital Mammography X-Ray Image IOD | Image | DX_Image | (0028,1054) | Rescale Type | Rescale​Type | 1 | LO | 1 | Specifies the output units of Rescale Slope (0028,1053) and Rescale Intercept (0028,1052). Enumerated Values: US Unspecified | Table C.8-70. DX Image Module Attributes |  | Required
Digital Mammography X-Ray Image IOD | Image | DX_Image | (2050,0020) | Presentation LUT Shape | Presentation​LUT​Shape | 1 | CS | 1 | Specifies an identity transformation for the Presentation LUT, other than to account for the Value of Photometric Interpretation (0028,0004), such that the output of all grayscale transformations defined in the IOD containing this Module are defined to be P-Values. Enumerated Values: IDENTITY, INVERSE | Table C.8-70. DX Image Module Attributes | {'Enumerated Values': ['IDENTITY', 'INVERSE']} | Required
Digital Mammography X-Ray Image IOD | Image | DX_Image | (0028,2110) | Lossy Image Compression | Lossy​Image​Compression | 1 | CS | 1 | Specifies whether an Image has undergone lossy compression (at a point in its lifetime). Enumerated Values: 00, 01 | Table C.8-70. DX Image Module Attributes | {'Enumerated Values': ['00', '01', 00, 01]} | Required
Digital Mammography X-Ray Image IOD | Image | DX_Image | (0028,2112) | Lossy Image Compression Ratio | Lossy​Image​Compression​Ratio | 1C | DS | 1-n | Describes the approximate lossy compression ratio(s) that have been applied to this image. Required if Lossy Image Compression (0028,2110) is '01'. | Table C.8-70. DX Image Module Attributes |  | Conditional
Digital Mammography X-Ray Image IOD | Image | DX_Image | (0008,2111) | Derivation Description | Derivation​Description | 3 | ST | 1 | A text description of how this image was derived. See Section C.8.11.3.1.4 for further explanation. | Table C.8-70. DX Image Module Attributes |  | Optional
Digital Mammography X-Ray Image IOD | Image | DX_Image | (0018,1400) | Acquisition Device Processing Description | Acquisition​Device​Processing​Description | 3 | LO | 1 | Indicates any visual processing performed on the images prior to exchange. See Section C.8.11.3.1.3 for further explanation. | Table C.8-70. DX Image Module Attributes |  | Optional
Digital Mammography X-Ray Image IOD | Image | DX_Image | (0018,1401) | Acquisition Device Processing Code | Acquisition​Device​Processing​Code | 3 | LO | 1 | Code representing the device-specific processing associated with the image (e.g., Organ Filtering code) | Table C.8-70. DX Image Module Attributes |  | Optional
Digital Mammography X-Ray Image IOD | Image | DX_Image | (0020,0020) | Patient Orientation | Patient​Orientation | 1C | CS | 2 | Patient direction of the rows and columns of the image. Required if View Code Sequence (0054,0220) is not present or is present with an Item value other than (119376003, SCT, 'tissue specimen') or (127457009, SCT, 'tissue specimen from breast'). May be present otherwise. | Table C.8-70. DX Image Module Attributes | {'Defined Terms': ["A", "P", "R", "L", "H", "F"]} | Conditional
Digital Mammography X-Ray Image IOD | Image | DX_Image | (0050,0004) | Calibration Image | Calibration​Image | 3 | CS | 1 | Indicates whether a reference object (phantom) of known size is present in the image and was used for calibration. Enumerated Values: YES, NO | Table C.8-70. DX Image Module Attributes | {'Enumerated Values': ['YES', 'NO']} | Optional
Digital Mammography X-Ray Image IOD | Image | DX_Image | (0028,0301) | Burned In Annotation | Burned​In​Annotation | 1 | CS | 1 | Indicates whether or not image contains sufficient burned in annotation to identify the patient and date the image was acquired. Enumerated Values: YES, NO | Table C.8-70. DX Image Module Attributes | {'Enumerated Values': ['YES', 'NO']} | Required
Digital Mammography X-Ray Image IOD | Image | DX_Image | (0028,3010) | VOI LUT Sequence | VOILUT​Sequence | 1C | SQ | 1 | Defines a Sequence of VOI LUTs. See Section C.8.11.3.1.5 for further explanation. | Table C.8-70. DX Image Module Attributes |  | Conditional
Digital Mammography X-Ray Image IOD | Image | DX_Image | (0028,1050) | Window Center | Window​Center | 1C | DS | 1-n | Defines a Window Center for display. See Section C.8.11.3.1.5 for further explanation. | Table C.8-70. DX Image Module Attributes |  | Conditional
Digital Mammography X-Ray Image IOD | Image | DX_Image | (0028,1051) | Window Width | Window​Width | 1C | DS | 1-n | Window Width for display. See Section C.8.11.3.1.5 for further explanation. | Table C.8-70. DX Image Module Attributes |  | Conditional
Digital Mammography X-Ray Image IOD | Image | DX_Image | (0028,1055) | Window Center & Width Explanation | Window​Center​Width​Explanation | 3 | LO | 1-n | Free form explanation of the meaning of the Window Center and Width. Multiple Values correspond to multiple Window Center and Width Values. | Table C.8-70. DX Image Module Attributes |  | Optional
Digital Mammography X-Ray Image IOD | Series | DX_Series | (0008,0060) | Modality | Modality | 1 | CS | 1 | Type of device, process or method that originally acquired the data used to create the images in this Series. | Table C.8-68. DX Series Module Attributes | {'Enumerated Values': ['DX', 'PX', 'IO', 'MG']} | Required
Digital Mammography X-Ray Image IOD | Series | DX_Series | (0008,1111) | Referenced Performed Procedure Step Sequence | Referenced​Performed​Procedure​Step​Sequence | 1C | SQ | 1 | Uniquely identifies the Performed Procedure Step SOP Instance to which the Series is related. Only a single Item shall be included in this Sequence. Required if a Performed Procedure Step SOP Class was involved in the creation of this Series. | Table C.8-68. DX Series Module Attributes |  | Conditional
Digital Mammography X-Ray Image IOD | Series | DX_Series | (0008,0068) | Presentation Intent Type | Presentation​Intent​Type | 1 | CS | 1 | Identifies the intent of the images that are contained within this Series. See Section C.8.11.1.1.1 for further explanation. | Table C.8-68. DX Series Module Attributes | {'Enumerated Values': ['FOR PRESENTATION', 'FOR PROCESSING']} | Required
Digital Mammography X-Ray Image IOD | Image | Mammography_Image | (0008,0008) | Image Type | Image​Type | 1 | CS | 2-n | {'Enumerated Values': ['']} | Table C.8-74. Mammography Image Module Attributes | {'Enumerated Values' : ['ORIGINAL ', 'DERIVED', 'PRIMARY ', 'SECONDARY', '']} | Required
Digital Mammography X-Ray Image IOD | Image | Mammography_Image | (0018,1508) | Positioner Type | Positioner​Type | 1 | CS | 1 | {'Enumerated Values': ['MAMMOGRAPHIC', 'NONE']} | Table C.8-74. Mammography Image Module Attributes | {'Enumerated Values': ['MAMMOGRAPHIC', 'NONE']} | Required
Digital Mammography X-Ray Image IOD | Image | Mammography_Image | (0018,1110) | Distance Source to Detector | Distance​Source​To​Detector | 3 | DS | 1 |  | Table C.8-74. Mammography Image Module Attributes |  | Optional
Digital Mammography X-Ray Image IOD | Image | Mammography_Image | (0018,1111) | Distance Source to Patient | Distance​Source​To​Patient | 3 | DS | 1 |  | Table C.8-74. Mammography Image Module Attributes |  | Optional
Digital Mammography X-Ray Image IOD | Image | Mammography_Image | (0018,1510) | Positioner Primary Angle | Positioner​Primary​Angle | 3 | DS | 1 |  | Table C.8-74. Mammography Image Module Attributes |  | Optional
Digital Mammography X-Ray Image IOD | Image | Mammography_Image | (0018,9559) | Positioner Primary Angle Direction | Positioner​Primary​Angle​Direction | 3 | CS | 1 | {'Enumerated Values': ['CW', 'CC']} | Table C.8-74. Mammography Image Module Attributes | {'Enumerated Values': ['CW', 'CC']} | Optional
Digital Mammography X-Ray Image IOD | Image | Mammography_Image | (0018,1511) | Positioner Secondary Angle | Positioner​Secondary​Angle | 3 | DS | 1 |  | Table C.8-74. Mammography Image Module Attributes |  | Optional
Digital Mammography X-Ray Image IOD | Image | Mammography_Image | (0020,0062) | Image Laterality | Image​Laterality | 1 | CS | 1 | {'Enumerated Values': ['R', 'L', 'B']} | Table C.8-74. Mammography Image Module Attributes | {'Enumerated Values': ['R', 'L', 'B']} | Required
Digital Mammography X-Ray Image IOD | Image | Mammography_Image | (0040,0318) | Organ Exposed | Organ​Exposed | 1 | CS | 1 | {'Enumerated Values': ['BREAST']} | Table C.8-74. Mammography Image Module Attributes | {'Enumerated Values': ['BREAST']} | Required
Digital Mammography X-Ray Image IOD | Image | Mammography_Image | (0028,1300) | Breast Implant Present | Breast​Implant​Present | 3 | CS | 1 | {'Enumerated Values': ['YES', 'NO']} | Table C.8-74. Mammography Image Module Attributes | {'Enumerated Values': ['YES', 'NO']} | Optional
Digital Mammography X-Ray Image IOD | Image | Mammography_Image | (0028,1350) | Partial View | Partial​View | 3 | CS | 1 | {'Enumerated Values': ['YES', 'NO']} | Table C.8-74. Mammography Image Module Attributes | {'Enumerated Values': ['YES', 'NO']} | Optional
Digital Mammography X-Ray Image IOD | Image | Mammography_Image | (0028,1351) | Partial View Description | Partial​View​Description | 3 | ST | 1 |  | Table C.8-74. Mammography Image Module Attributes |  | Optional
Digital Mammography X-Ray Image IOD | Image | Mammography_Image | (0028,1352) | Partial View Code Sequence | Partial​View​Code​Sequence | 3 | SQ | 1 |  | Table C.8-74. Mammography Image Module Attributes |  | Optional
Digital Mammography X-Ray Image IOD | Image | Mammography_Image | (0008,2218) | Anatomic Region Sequence | Anatomic​Region​Sequence | 1 | SQ | 1 | Sequence that identifies the anatomic region of interest in this Instance (i.e., external anatomy, surface anatomy, or general region of the body). Zero or one Item shall be included in this Sequence. | Table 10-5. General Anatomy Mandatory Macro Attributes |  | Required
Digital Mammography X-Ray Image IOD | Image | Mammography_Image | (0008,2220) | Anatomic Region Modifier Sequence | Anatomic​Region​Modifier​Sequence | 3 | SQ | 1 | Sequence of Items that modifies the anatomic region of interest of this Instance. One or more Items are permitted in this Sequence. | Table 10-5. General Anatomy Mandatory Macro Attributes |  | Optional
Digital Mammography X-Ray Image IOD | Image | Mammography_Image | (0008,2228) | Primary Anatomic Structure Sequence | Primary​Anatomic​Structure​Sequence | 3 | SQ | 1 | Sequence of Items that identifies the primary anatomic structure(s) of interest in this Instance. One or more Items are permitted in this Sequence. | Table 10-8. Primary Anatomic Structure Macro Attributes |  | Optional
Digital Mammography X-Ray Image IOD | Image | Mammography_Image | (0054,0220) | View Code Sequence | View​Code​Sequence | 1 | SQ | 1 | CID 4014 “View for Mammography”. | Table C.8-74. Mammography Image Module Attributes |  | Required
Digital Mammography X-Ray Image IOD | Image | Mammography_Image | (0054,0222) | View Modifier Code Sequence | View​Modifier​Code​Sequence | 2 | SQ | 1 | CID 4015 “View Modifier for Mammography”. | Table C.8-74. Mammography Image Module Attributes |  | Required
Digital Mammography X-Ray Image IOD | Image | Mammography_Image | (0018,2041) | Biopsy Target Sequence | Biopsy​Target​Sequence | 3 | SQ | 1 | Sequence that identifies the targets of a biopsy procedure. One or more Items are permitted in this Sequence, where the coordinate system is the same for all Items within a pair of stereo images. | Table C.8-74. Mammography Image Module Attributes |  | Optional
Digital Mammography X-Ray Image IOD | Series | Mammography_Series | (0008,0060) | Modality | Modality | 1 | CS | 1 | Type of device, process or method that originally acquired the data used to create the images in this Series. | Table C.8-73. Mammography Series Module Attributes | {'Enumerated Values': ['MG']} | Required
Digital Mammography X-Ray Image IOD | Series | Mammography_Series | (0040,0275) | Request Attributes Sequence | Request​Attributes​Sequence | 3 | SQ | 1 | Sequence that contains Attributes from the Imaging Service Request. One or more Items are permitted in this Sequence. | Table C.8-73. Mammography Series Module Attributes |  | Optional
MR Image IOD | Image | MR_Image | (0008,0008) | Image Type | Image​Type | 1 | CS | 2-n | Image identification characteristics. See Section C.8.3.1.1.1 for specialization. | Table C.8-4. MR Image Module Attributes | {'Enumerated Values' : ['ORIGINAL ', 'DERIVED', 'PRIMARY ', 'SECONDARY', 'DENSITY MAP', 'DIFFUSION MAP', 'IMAGE ADDITION','MODULUS SUBTRACT','MPR','OTHER', 'PHASE MAP','PHASE SUBTRACT','PROJECTION IMAGE','T1 MAP','T2 MAP','VELOCITY MAP']} | Required
MR Image IOD | Image | MR_Image | (0028,0002) | Samples per Pixel | Samples​Per​Pixel | 1 | US | 1 | Number of samples (planes) in this image. See Section C.8.3.1.1.2 for specialization. | Table C.8-4. MR Image Module Attributes |  | Required
MR Image IOD | Image | MR_Image | (0028,0004) | Photometric Interpretation | Photometric​Interpretation | 1 | CS | 1 | Specifies the intended interpretation of the pixel data. See Section C.8.3.1.1.3 for specialization. | Table C.8-4. MR Image Module Attributes | {'Enumerated Values': ['MONOCHROME1', 'MONOCHROME2']} | Required
MR Image IOD | Image | MR_Image | (0028,0100) | Bits Allocated | Bits​Allocated | 1 | US | 1 | Number of bits allocated for each pixel sample. Each sample shall have the same number of bits allocated. See Section C.8.3.1.1.4 for specialization. | Table C.8-4. MR Image Module Attributes |  | Required
MR Image IOD | Image | MR_Image | (0028,0101) | Bits Stored | Bits​Stored | 1 | US | 1 | Number of bits stored for each pixel sample. Each sample shall have the same number of bits stored. | Table C.8-4. MR Image Module Attributes |  | Required
MR Image IOD | Image | MR_Image | (0028,0102) | High Bit | High​Bit | 1 | US | 1 | Most significant bit for pixel sample data. Each sample shall have the same high bit. High Bit (0028,0102) shall be one less than Bits Stored (0028,0101). | Table C.8-4. MR Image Module Attributes |  | Required
MR Image IOD | Image | MR_Image | (0018,0020) | Scanning Sequence | Scanning​Sequence | 1 | CS | 1-n | Description of the type of data taken. | Table C.8-4. MR Image Module Attributes | {'Enumerated Values': ['SE', 'IR', 'GR', 'EP', 'RM']} | Required
MR Image IOD | Image | MR_Image | (0018,0021) | Sequence Variant | Sequence​Variant | 1 | CS | 1-n | Variant of the Scanning Sequence. | Table C.8-4. MR Image Module Attributes | {'Defined Terms': ['SK', 'MTC', 'SS', 'TRSS', 'SP', 'MP', 'OSP', 'NONE']} | Required
MR Image IOD | Image | MR_Image | (0018,0022) | Scan Options | Scan​Options | 2 | CS | 1-n | Parameters of scanning sequence. | Table C.8-4. MR Image Module Attributes | {'Defined Terms': ['PER', 'RG', 'CG', 'PPG', 'FC', 'PFF', 'PFP', 'SP', 'FS']} | Required
MR Image IOD | Image | MR_Image | (0018,0023) | MR Acquisition Type | MR​Acquisition​Type | 2 | CS | 1 | Identification of data encoding scheme. | Table C.8-4. MR Image Module Attributes | {'Enumerated Values': ['2D', '3D']} | Required
MR Image IOD | Image | MR_Image | (0018,0080) | Repetition Time | Repetition​Time | 2C | DS | 1 | The period of time in msec between the beginning of a pulse sequence and the beginning of the succeeding (essentially identical) pulse sequence. | Table C.8-4. MR Image Module Attributes |  | Conditional
MR Image IOD | Image | MR_Image | (0018,0081) | Echo Time | Echo​Time | 2 | DS | 1 | Time in ms between the middle of the excitation pulse and the peak of the echo produced (kx=0). | Table C.8-4. MR Image Module Attributes |  | Required
MR Image IOD | Image | MR_Image | (0018,0091) | Echo Train Length | Echo​Train​Length | 2 | IS | 1 | Number of lines in k-space acquired per excitation per image. | Table C.8-4. MR Image Module Attributes |  | Required
MR Image IOD | Image | MR_Image | (0018,0082) | Inversion Time | Inversion​Time | 2C | DS | 1 | Time in msec after the middle of inverting RF pulse to middle of excitation pulse to detect the amount of longitudinal magnetization. | Table C.8-4. MR Image Module Attributes |  | Conditional
MR Image IOD | Image | MR_Image | (0018,1060) | Trigger Time | Trigger​Time | 2C | DS | 1 | Time, in msec, between peak of the R wave and the peak of the echo produced. | Table C.8-4. MR Image Module Attributes |  | Conditional
MR Image IOD | Image | MR_Image | (0018,0024) | Sequence Name | Sequence​Name | 3 | SH | 1 | User defined name for the combination of Scanning Sequence (0018,0020) and Sequence Variant (0018,0021). | Table C.8-4. MR Image Module Attributes |  | Optional
MR Image IOD | Image | MR_Image | (0018,0025) | Angio Flag | Angio​Flag | 3 | CS | 1 | Angio Image Indicator. Primary image for Angio processing. | Table C.8-4. MR Image Module Attributes | {'Enumerated Values': ['Y', 'N']} | Optional
MR Image IOD | Image | MR_Image | (0018,0083) | Number of Averages | Number​Of​Averages | 3 | DS | 1 | Number of times a given pulse sequence is repeated before any parameter is changed. | Table C.8-4. MR Image Module Attributes |  | Optional
MR Image IOD | Image | MR_Image | (0018,0084) | Imaging Frequency | Imaging​Frequency | 3 | DS | 1 | Precession frequency in MHz of the nucleus being addressed. | Table C.8-4. MR Image Module Attributes |  | Optional
MR Image IOD | Image | MR_Image | (0018,0085) | Imaged Nucleus | Imaged​Nucleus | 3 | SH | 1 | Nucleus that is resonant at the imaging frequency. Examples: 31P, 1H. | Table C.8-4. MR Image Module Attributes |  | Optional
MR Image IOD | Image | MR_Image | (0018,0086) | Echo Number(s) | Echo​Numbers | 3 | IS | 1-n | The echo number used in generating this image. | Table C.8-4. MR Image Module Attributes |  | Optional
MR Image IOD | Image | MR_Image | (0018,0087) | Magnetic Field Strength | Magnetic​Field​Strength | 3 | DS | 1 | Nominal field strength of MR magnet, in Tesla. | Table C.8-4. MR Image Module Attributes |  | Optional
MR Image IOD | Image | MR_Image | (0018,0089) | Number of Phase Encoding Steps | Number​Of​Phase​Encoding​Steps | 3 | IS | 1 | Total number of lines in k-space in the 'y' direction collected during acquisition. | Table C.8-4. MR Image Module Attributes |  | Optional
MR Image IOD | Image | MR_Image | (0018,0093) | Percent Sampling | Percent​Sampling | 3 | DS | 1 | Fraction of acquisition matrix lines acquired, expressed as a percent. | Table C.8-4. MR Image Module Attributes |  | Optional
MR Image IOD | Image | MR_Image | (0018,0094) | Percent Phase Field of View | Percent​Phase​Field​Of​View | 3 | DS | 1 | Ratio of field of view dimension in phase direction to field of view dimension in frequency direction, expressed as a percent. | Table C.8-4. MR Image Module Attributes |  | Optional
MR Image IOD | Image | MR_Image | (0018,0095) | Pixel Bandwidth | Pixel​Bandwidth | 3 | DS | 1 | Reciprocal of the total sampling period, in hertz per pixel. | Table C.8-4. MR Image Module Attributes |  | Optional
MR Image IOD | Image | MR_Image | (0018,1062) | Nominal Interval | Nominal​Interval | 3 | IS | 1 | Average R-R interval used for the scans, in msec. | Table C.8-4. MR Image Module Attributes |  | Optional
MR Image IOD | Image | MR_Image | (0018,1080) | Beat Rejection Flag | Beat​Rejection​Flag | 3 | CS | 1 | Beat length sorting has been applied. | Table C.8-4. MR Image Module Attributes | {'Enumerated Values': ['Y', 'N']} | Optional
MR Image IOD | Image | MR_Image | (0018,1081) | Low R-R Value | Low​RR​Value | 3 | IS | 1 | R-R interval low limit for beat rejection, in msec. | Table C.8-4. MR Image Module Attributes |  | Optional
MR Image IOD | Image | MR_Image | (0018,1082) | High R-R Value | High​RR​Value | 3 | IS | 1 | R-R interval high limit for beat rejection, in msec. | Table C.8-4. MR Image Module Attributes |  | Optional
MR Image IOD | Image | MR_Image | (0018,1083) | Intervals Acquired | Intervals​Acquired | 3 | IS | 1 | Number of R-R intervals acquired. | Table C.8-4. MR Image Module Attributes |  | Optional
MR Image IOD | Image | MR_Image | (0018,1084) | Intervals Rejected | Intervals​Rejected | 3 | IS | 1 | Number of R-R intervals rejected. | Table C.8-4. MR Image Module Attributes |  | Optional
MR Image IOD | Image | MR_Image | (0018,1085) | PVC Rejection | PVC​Rejection | 3 | LO | 1 | Description of type of PVC rejection criteria used. | Table C.8-4. MR Image Module Attributes |  | Optional
MR Image IOD | Image | MR_Image | (0018,1086) | Skip Beats | Skip​Beats | 3 | IS | 1 | Number of beats skipped after a detected arrhythmia. | Table C.8-4. MR Image Module Attributes |  | Optional
MR Image IOD | Image | MR_Image | (0018,1088) | Heart Rate | Heart​Rate | 3 | IS | 1 | Beats per minute. | Table C.8-4. MR Image Module Attributes |  | Optional
MR Image IOD | Image | MR_Image | (0018,1090) | Cardiac Number of Images | Cardiac​Number​Of​Images | 3 | IS | 1 | Number of images per cardiac cycle. | Table C.8-4. MR Image Module Attributes |  | Optional
MR Image IOD | Image | MR_Image | (0018,1094) | Trigger Window | Trigger​Window | 3 | IS | 1 | Percent of R-R interval, based on Heart Rate (0018,1088), prescribed as a window for a valid/usable trigger. | Table C.8-4. MR Image Module Attributes |  | Optional
MR Image IOD | Image | MR_Image | (0018,1100) | Reconstruction Diameter | Reconstruction​Diameter | 3 | DS | 1 | Diameter in mm. of the region from within which data were used in creating the reconstruction of the image. | Table C.8-4. MR Image Module Attributes |  | Optional
MR Image IOD | Image | MR_Image | (0018,1250) | Receive Coil Name | Receive​Coil​Name | 3 | SH | 1 | Receive coil used. | Table C.8-4. MR Image Module Attributes |  | Optional
MR Image IOD | Image | MR_Image | (0018,1251) | Transmit Coil Name | Transmit​Coil​Name | 3 | SH | 1 | Transmit coil used. | Table C.8-4. MR Image Module Attributes |  | Optional
MR Image IOD | Image | MR_Image | (0018,1310) | Acquisition Matrix | Acquisition​Matrix | 3 | US | 4 | Dimensions of the acquired frequency /phase data before reconstruction. | Table C.8-4. MR Image Module Attributes |  | Optional
MR Image IOD | Image | MR_Image | (0018,1312) | In-plane Phase Encoding Direction | In​Plane​Phase​Encoding​Direction | 3 | CS | 1 | The axis of phase encoding with respect to the image. | Table C.8-4. MR Image Module Attributes | {'Enumerated Values': ['ROW', 'COL']} | Optional
MR Image IOD | Image | MR_Image | (0018,1314) | Flip Angle | Flip​Angle | 3 | DS | 1 | Steady state angle in degrees to which the magnetic vector is flipped from the magnetic vector of the primary field. | Table C.8-4. MR Image Module Attributes |  | Optional
MR Image IOD | Image | MR_Image | (0018,1316) | SAR | SAR | 3 | DS | 1 | Calculated whole body Specific Absorption Rate in watts/kilogram. | Table C.8-4. MR Image Module Attributes |  | Optional
MR Image IOD | Image | MR_Image | (0018,1315) | Variable Flip Angle Flag | Variable​Flip​Angle​Flag | 3 | CS | 1 | Flip angle variation applied during image acquisition. | Table C.8-4. MR Image Module Attributes | {'Enumerated Values': ['Y', 'N']} | Optional
MR Image IOD | Image | MR_Image | (0018,1318) | dB/dt | d​Bdt | 3 | DS | 1 | The rate of change of the gradient coil magnetic flux density with time (T/s). | Table C.8-4. MR Image Module Attributes |  | Optional
MR Image IOD | Image | MR_Image | (0020,0100) | Temporal Position Identifier | Temporal​Position​Identifier | 3 | IS | 1 | Temporal order of a dynamic or functional set of Images. | Table C.8-4. MR Image Module Attributes |  | Optional
MR Image IOD | Image | MR_Image | (0020,0105) | Number of Temporal Positions | Number​Of​Temporal​Positions | 3 | IS | 1 | Total number of temporal positions prescribed. | Table C.8-4. MR Image Module Attributes |  | Optional
MR Image IOD | Image | MR_Image | (0020,0110) | Temporal Resolution | Temporal​Resolution | 3 | DS | 1 | Time delta between Images in a dynamic or functional set of Images. | Table C.8-4. MR Image Module Attributes |  | Optional
MR Image IOD | Image | MR_Image | (0018,1320) | B1rms | B1rms | 3 | FL | 1 | B1+ rms value in units of microtesla (µT) for the acquisition producing the image. | Table C.8-4. MR Image Module Attributes |  | Optional
MR Image IOD | Image | MR_Image | (0008,2218) | Anatomic Region Sequence | Anatomic​Region​Sequence | 3 | SQ | 1 | Sequence that identifies the anatomic region of interest in this Instance (i.e., external anatomy, surface anatomy, or general region of the body). Only a single Item is permitted in this Sequence. | Table 10-7. General Anatomy Optional Macro Attributes |  | Optional
MR Image IOD | Image | MR_Image | (0008,2228) | Primary Anatomic Structure Sequence | Primary​Anatomic​Structure​Sequence | 3 | SQ | 1 | Sequence of Items that identifies the primary anatomic structure(s) of interest in this Instance. One or more Items are permitted in this Sequence. | Table 10-8. Primary Anatomic Structure Macro Attributes |  | Optional
MR Image IOD | Image | MR_Image | (0054,0220) | View Code Sequence | View​Code​Sequence | 3 | SQ | 1 | Sequence that describes the projection of the anatomic region of interest. Only a single Item is permitted in this Sequence. | Table 10-25. Optional View and Slice Progression Direction Macro Attributes |  | Optional
MR Image IOD | Image | MR_Image | (0054,0500) | Slice Progression Direction | Slice​Progression​Direction | 3 | CS | 1 | Describes the anatomical direction in which a set of slices is progressing (see Section 10.20.1.1). Meaningful only for cardiac images. Enumerated Values are defined in Section 10.20.1.1. | Table 10-25. Optional View and Slice Progression Direction Macro Attributes | {'Enumerated Values': ['APEX_TO_BASE', 'BASE_TO_APEX', 'ANT_TO_INF', 'INF_TO_ANT', 'SEPTUM_TO_WALL', 'WALL_TO_SEPTUM']} | Optional
MR Image IOD | Image | MR_Image | (300A,0675) | Equipment Frame of Reference UID | Equipment​Frame​Of​Reference​UID | 1C | UI | 1 | Frame of Reference identifier for the Treatment Delivery Device. Required if either Imaging Equipment to Treatment Delivery Device Relationship Sequence (300A,07A1) or Patient to Equipment Relationship Sequence (300A,07A0) is present. See Section C.36.12.1 and Section C.36.2.4.12.1.1. | Table C.36.2.4.12-1 RT Equipment Mapping and Plan Reference Macro Attributes |  | Conditional
MR Image IOD | Image | MR_Image | (300A,07A1) | Imaging Equipment to Treatment Delivery Device Relationship Sequence | Imaging​Equipment​To​Treatment​Delivery​Device​Relationship​Sequence | 3 | SQ | 1 | The relationship between the coordinate system of the Imaging Device and coordinate system of the Treatment Delivery Device. Only a single Item is permitted in this Sequence. See Section 10.39.1.1. | Table C.36.2.4.12-1 RT Equipment Mapping and Plan Reference Macro Attributes |  | Optional
MR Image IOD | Image | MR_Image | (300A,012C) | Isocenter Position | Isocenter​Position | 3 | DS | 3 | Isocenter coordinates (x,y,z), in mm. Specifies the location of the machine isocenter in the Patient-Based Coordinate System associated with the Frame of Reference. It allows transformation from the Equipment-Based Coordinate System to the Patient-Based Coordinate System. | Table C.36.2.4.12-1 RT Equipment Mapping and Plan Reference Macro Attributes |  | Optional
MR Image IOD | Image | MR_Image | (300A,07A0) | Patient to Equipment Relationship Sequence | Patient​To​Equipment​Relationship​Sequence | 3 | SQ | 1 | The relationship between the Patient Frame of Reference and the coordinate system of the Treatment Delivery Device. Only a single Item is permitted in this Sequence. | Table C.36.2.4.12-1 RT Equipment Mapping and Plan Reference Macro Attributes |  | Optional
MR Image IOD | Image | MR_Image | (300A,079F) | Patient Treatment Preparation Sequence | Patient​Treatment​Preparation​Sequence | 3 | SQ | 1 | Procedures and devices required to prepare a patient for treatment. Only a single Item is permitted in this Sequence. See Section C.8.8.12.1.3. | Table C.36.2.4.12-1 RT Equipment Mapping and Plan Reference Macro Attributes |  | Optional
MR Image IOD | Image | MR_Image | (300C,0002) | Referenced RT Plan Sequence | Referenced​RT​Plan​Sequence | 3 | SQ | 1 | RT Plan Instance associated with the acquisition of this Instance. Only a single Item is permitted in this Sequence. | Table C.36.2.4.12-1 RT Equipment Mapping and Plan Reference Macro Attributes |  | Optional
Computed Radiography Image IOD | Series | CR_Series | (0018,0015) | Body Part Examined | Body​Part​Examined | 2 | CS | 1 | {} | Table C.8-1. CR Series Module Attributes | {'Defined Terms': ['BREAST', 'CHEST', 'SSPINE']} | Required
Computed Radiography Image IOD | Series | CR_Series | (0018,5101) | View Position | View​Position | 2 | CS | 1 | {'Defined Terms': ['AP', 'PA', 'LL', 'RL', 'RLD', 'LLD', 'RLO', 'LLO']} | Table C.8-1. CR Series Module Attributes | {'Defined Terms': ['AP', 'PA', 'LL', 'RL', 'RLD', 'LLD', 'RLO', 'LLO']} | Required
Computed Radiography Image IOD | Series | CR_Series | (0018,1160) | Filter Type | Filter​Type | 3 | SH | 1 | {} | Table C.8-1. CR Series Module Attributes |  | Optional
Computed Radiography Image IOD | Series | CR_Series | (0018,1180) | Collimator/grid Name | Collimator​Grid​Name | 3 | SH | 1 | {} | Table C.8-1. CR Series Module Attributes |  | Optional
Computed Radiography Image IOD | Series | CR_Series | (0018,1190) | Focal Spot(s) | Focal​Spots | 3 | DS | 1-n | {} | Table C.8-1. CR Series Module Attributes |  | Optional
Computed Radiography Image IOD | Series | CR_Series | (0018,1260) | Plate Type | Plate​Type | 3 | SH | 1 | {} | Table C.8-1. CR Series Module Attributes |  | Optional
Computed Radiography Image IOD | Series | CR_Series | (0018,1261) | Phosphor Type | Phosphor​Type | 3 | LO | 1 | {} | Table C.8-1. CR Series Module Attributes |  | Optional